Mt. Morgan SITE WORK INSTRUCTION

Slurry Viscosity and Rheology Test

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-204-008 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-204-008 |
| Area | 204 Met Lab - Physical Testwork |
| Parent procedure | HM-PRC-VXX-PRO208 Metallurgical Laboratory Physical Testwork |
| Classification | Operating |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

DRAFT FOR REVIEW. Martabe work instruction 'Measure Viscosity Using Brookfield DV2TLV Viscometer' (DOC-3-MET-PMC-WIN-00128-IE, v1.0, issued 25/12/2024) added as Section 6 Part B. Not approved until reviewed. JSEA JSEA-PRO-MET-204-008 must be checked.

# 1  TASK AND SCOPE

Measure the viscosity and flow behaviour of a slurry at the operating density.

Viscosity governs pumping duty, mixing and thickener behaviour, and rises sharply with fines.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: On each new ore type, and whenever pumping or mixing behaviour changes.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO208 - Metallurgical Laboratory Physical Testwork. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-204-008 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-204-008.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-204-008 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Splash of cyanide-bearing slurry or solution to skin or eyes | Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. |
| Contact with an agitator, sample cutter, pump or screen that starts without warning | Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |
| Manual handling injury from carrying sample buckets, containers and composite drums | Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. |
| Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. |
| [JSEA to confirm] Sample contact and splash to the eyes at the viscometer | Latex gloves, glasses and safety boots. Eyewash proven flowing; wash contact with soap under running water. Handle the sample carefully. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-204-008 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- Viscometer

- Sample container and stand

- Thermometer

- Marcy scale

- Stopwatch

- [M] Brookfield DV2TLV viscometer with spindles and its measuring cup, clean water for cleaning

# 6  STEPS

Part A - Slurry viscosity and rheology

1.  Prepare the sample at the density the circuit actually runs at, not at a convenient density.

2.  Record the density and temperature - both change the result substantially.

3.  Homogenise the sample immediately before measuring; a settled slurry reads low.

4.  Measure viscosity at the range of shear rates the method calls for, not a single point. Mineral slurries are not Newtonian.

5.  Record the full shear rate against shear stress series.

6.  Repeat at two other densities where the test is being used to size equipment.

7.  Report the rheology curve, the density and the temperature together. A viscosity quoted without them means nothing.

Part B - Operating the Brookfield DV2TLV viscometer [M]

[M] = taken from the Martabe team work instruction "Measure Viscosity Using Brookfield DV2TLV Viscometer (DOC-3-MET-PMC-WIN-00128-IE)" (25/12/2024). [CONFIRM] = unspecified or inconsistent in the source; resolve before this instruction is approved for use. Part A (original steps) is unchanged.

[CONFIRM] the source reads one viscosity point at about 20 rpm for daily work. Part A requires a shear rate series, and the density and temperature to be recorded with the result; apply Part A on top of these steps. The instrument manual and the source's 'Brookfield DV2T operating instruction' are not received.

8.  Take 5 and prepare the viscometer, spindles and measuring cup; check they work. Have clean water ready [M].

9.  Make sure the viscometer is level. Switch it on with the power button.

10.  Autozero by pressing 'next' on the screen. The spindle must not be fitted during autozero.

11.  Choose the spindle. Each spindle has a minimum and maximum viscosity it can measure; see the range in the manual. The spindle and speed chosen must be able to measure all the samples, and all samples must be measured with the same spindle and speed.

12.  Fit the spindle. In the settings menu choose 'configure viscosity test'. Set the speed between 1 and 200 rpm (20 rpm for daily work) and the measuring time for each sample.

13.  Pour the sample into the cup. Bring the sample surface level with the mark on the spindle. Lower the spindle into the sample by adjusting the viscometer position.

14.  Press run and record the viscosity reading.

15.  If the torque is below 10 per cent or above 90 per cent, change the speed and the spindle. A slower speed and a smaller spindle can measure higher viscosity.

16.  Clean-up: switch the viscometer off, take the spindle off, clean it with clean water and put it back in its box.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- KBK-MIR-MP-PRO-OPE-SOP-0016 Uji Viskositas

- Martabe work instruction 'Measure Viscosity Using Brookfield DV2TLV Viscometer' DOC-3-MET-PMC-WIN-00128-IE v1.0 (issued 25/12/2024; reference SOP DOC-IV-MET-CHH-SOP-00039) - Part B source