Safe Work Instruction

Balance Verification and Calibration Check

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, acid-resistant and nitrile rubber gloves, and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-206-003 - Balance Verification and Calibration Check |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-206-003 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Verify the laboratory balances against certified weights so every mass is defensible. Every assay, every moisture and every solution standard depends on a balance reading truly. Area: 206 Met Lab - Facility, Safety and Quality. Points and equipment covered: Met lab bench - physical testwork bench and instrument Frequency: Daily verification before use; full calibration at the external interval. This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO210 - Metallurgical Laboratory Facility, Safety and Quality. Read that procedure before carrying out this task for the first time. Register line: SWI-PRO-MET-206-003 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-206-003. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. Cuts from broken glassware, and burns from hot glassware — Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. Non-representative sample leading to a wrong process decision — Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-206-003 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-206-003 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - Class F1 weight set 10 mg to 200 g - GR-200 0.1 mg analytical balance - BM-22 micro balance - 150 kg platform balance - Draught shield - Balance log - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Balance Set-up - Balance not level or in a draught - Balance not warmed up |  | - Confirm the balance is level, clean and in a draught-free position before use. - Allow the balance to warm up for the period the manual specifies. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Verification with Certified Weights - Weight handled by hand - Reading outside tolerance not acted on |  | - Zero the balance, then verify with at least a low, a mid and a high certified weight from the F1 set. - Handle certified weights only with tongs or gloves. A fingerprint on a 10 mg weight is a real error. - Record each reading against the certified value and the allowable tolerance. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Out of Tolerance and Records - Out-of-tolerance balance used for rough work - External calibration missed |  | - Where any point is outside tolerance, take the balance out of service and label it. Do not use it for rough work in the meantime. - Record every verification in the balance log against the balance identity. - Arrange external calibration at the set interval and file the certificate. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| HM-PRC-VXX-PRO210 | Metallurgical Laboratory Facility, Safety and Quality |
| JSEA-PRO-MET-206-003 | JSEA — Balance Verification and Calibration Check |
| HM-MMM-WHS-REG-XXX-R00-SWI-SWMS | SWI / SWMS Register — tab SWI-MET |
| KBK-MIR-LAB-PRE-SOP-012 and WET-SOP-013, -014, -015 | Balance verification |
| ILS quotation Q0023979 | Class F1 weight set 10 mg to 200 g, GR-200 and BM-22 balances, 150 kg platform balance |

In the event of an emergency

1. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 2. Skin or Eye Contact with Slurry Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 3. Cuts from Broken Glassware Apply pressure to stop the bleeding, get a trained first aider, and report to the Shift Supervisor. Call CH19 “EMERGENCY” and 000 for a serious cut.