Safe Work Instruction

Ductless Fume Cabinet Filter Change and Breakthrough Check

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, acid-resistant and nitrile rubber gloves, and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-206-013 - Ductless Fume Cabinet Filter Change and Breakthrough Check |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-206-013 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Change the carbon filters on the ductless fume cabinet and check for breakthrough before it happens. A ductless cabinet with a saturated filter discharges everything it has captured straight back into the room, and gives no indication that it is doing so. Area: 206 Met Lab - Facility, Safety and Quality. Points and equipment covered: Met lab bench - physical testwork bench and instrument Frequency: Breakthrough check monthly; filter change on indication or at the manufacturer's interval. This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO210 - Metallurgical Laboratory Facility, Safety and Quality. Read that procedure before carrying out this task for the first time. Register line: SWI-PRO-MET-206-013 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-206-013. This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Inhalation of acid fume or reagent vapour at the bench — Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Contact with sulphuric or hydrochloric acid - burn to skin or eyes — Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. Manual handling injury from carrying sample buckets, containers and composite drums — Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-206-013 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-206-013 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm a second person is in the laboratory for any cyanide work. Cyanide work is not done alone. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - Purair 30-PP ductless cabinet - Replacement carbon filters - Airflow indicator - Detector tubes or photoionisation detector for breakthrough - Gloves and disposal bags - Filter log - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Airflow and Breakthrough Check - Spent filter discharging fume into the room - Airflow out of range |  | - Confirm the cabinet airflow is within range before any use, and record it. - Carry out the breakthrough check at the interval set - sample the discharge air with detector tubes or a suitable detector while a representative solvent or reagent is in use. - Any detection in the discharge means the filter is spent. Take the cabinet out of service immediately and label it. - Track filter life by usage, not by calendar alone. A cabinet used heavily saturates faster. CAUTION: Any detection in the discharge means the filter is spent. Take the cabinet out of service. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Filter Change - Contact with a spent filter - Seal not made on the new filter - Manual handling of the filter |  | - To change: isolate the cabinet, wear gloves, remove the spent filter and bag it immediately as chemical waste. - Fit the new filter, confirm the seal and re-check the airflow. - Record the change date, the filter batch and the airflow reading in the filter log. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Use Limits - Reagent the filter is not rated for |  | - Do not use the cabinet for any reagent the filter is not rated for. Carbon does not capture everything. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| HM-PRC-VXX-PRO210 | Metallurgical Laboratory Facility, Safety and Quality |
| JSEA-PRO-MET-206-013 | JSEA — Ductless Fume Cabinet Filter Change and Breakthrough Check |
| HM-MMM-WHS-REG-XXX-R00-SWI-SWMS | SWI / SWMS Register — tab SWI-MET |
| ILS quotation Q0023979 | Purair 30-PP ductless cabinet with carbon filters |
| Note | No procedure was held for this cabinet before this revision |

In the event of an emergency

1. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 2. Skin or Eye Contact with Reagent Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 3. HCl Gas or Acid Contact Leave the area upwind. Flush skin or eyes with copious water at the safety shower or eyewash for a minimum of 15 minutes. If breathing is affected call CH19 “EMERGENCY” and call 000. 4. Slip, Fall or Lifting Injury Call CH19 “EMERGENCY”, call 000 if the injury is serious, and notify the Shift Supervisor.