Safe Work Procedure

Metallurgical Laboratory Facility, Safety and Quality

SWP-210

| DOCUMENT HISTORY AND STATUS |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Document: | Metallurgical Laboratory Facility, Safety and Quality Safe Work Procedure |  |  | Version No.: | A |  |  |
| Document No: | SWP-210 |  |  | Owner: | Process Operations Superintendent |  |  |
| Issue Date: | 11/08/2026 |  |  | Date Reviewed: | - |  |  |
| History: | Process owner: Heritage Mineral Operations, Business unit Mt. Morgan. Document approver: Process Manager. |  |  |  |  |  |  |
| Date Reviewed: | Version No.: | Written by: |  | Reviewed by: |  | Authorised by: |  |
| 11/08/2026 | A | GGT (cost code 4034-033) |  |  |  |  |  |
| Revision Description: |  |  |  |  |  |  |  |
| Changes: |  |  |  |  |  |  |  |
| Rev A - first issue. Converted from the GGT Standard Work Procedure layout to the Heritage Safe Work Procedure format; content unchanged. |  |  |  |  |  |  |  |
| Action: | Name: |  | Position: |  |  |  | Date: |
| Prepared by | GGT (cost code 4034-033) |  | Document preparer |  |  |  | 11/08/2026 |
| Reviewed by |  |  | Process Operations Superintendent |  |  |  |  |
| Reviewed by |  |  |  |  |  |  |  |
| Authorised by |  |  | Process Manager |  |  |  |  |

This document is controlled electronically as part of Heritage Minerals Pty Limited’s Health & Safety Management System.  Please refer to Heritage Minerals Controlled Document Register for the current version.

Table of Contents

1.	Purpose and Scope	3

2.	Responsibilities	3

3.	Abbreviations and Definitions	3

4.	Potential Hazards	4

5.	Personal Protective Equipment	4

6.	Procedure	5

6.1	Introduction	5

6.2	Procedure and Implementation Steps	5

6.3	Schedule	6

6.4	Flowchart	6

7.	Training	6

8.	Document Traceability	7

9.	Review Criteria	7

10.	Legislation & References	7

10.1	Legislation	7

10.2	Source Documents	7

# Purpose and Scope

This procedure covers metallurgical laboratory facility, safety and quality at the Mt. Morgan processing plant.

Metallurgical laboratory facility, safety and quality - chemical receipt and storage, fume cupboard and scrubber operation, balance and meter calibration, cyanide and acid spill response, waste disposal, glassware, equipment operation, emergency response, quality control, and laboratory equipment commissioning.

Document details:

Hazard analyses beneath: 14 (JSEA-PRO-MET-206-001 to -014)

Work instructions beneath: 13

Process owner: Heritage Mineral Operations, Business unit Mt. Morgan

Document owner: Process Operations Superintendent

Document approver: Process Manager

# Responsibilities

| Staff | Responsibilities |
| --- | --- |
| Operator / Metallurgical Technician / Met Lab Technician | - Carry out the task to the schedule and to the work instruction. - Only carry out work where competent and authorised to do so. - Record what was actually done. Work not done is recorded as not done - never invented and never back-filled. - Stop and report if any control in the JSEA is not in place. |
| Shift Supervisor / Coordinator | - Authorise the task and confirm the area is fit before it starts. - Know that the task is in progress and where the person is. - Understand that much of the plant is automated and equipment may start or stop without warning. - Act on any off-specification result reported from the task. |
| Metallurgist | - Set and maintain the schedule - what is done, how often, and what is reported. - Review the results and act on the trends. - Investigate any variance outside the agreed tolerance before a period is closed. |
| Process Manager / Superintendent | - Ensure all personnel are trained and competent. - Ensure this procedure and the documents beneath it are reviewed at the stated interval. - Approve the risk ranking in the JSEAs. |

# Abbreviations and Definitions

| Term | Meaning |
| --- | --- |
| JSEA | Job safety and environmental analysis - one per task, giving the step-by-step hazards, controls and risk ranking. |
| SWI | Safe work instruction - one per task, giving the step-by-step method. |
| SWP | Safe work procedure - this document. |
| Steady state | Normal operation held for at least 15 to 20 minutes. |
| ReCYN | The resin-in-pulp cyanide and metal recovery process in Areas 041, 051, 071 and 081. |
| CN Free / WAD CN | Free cyanide by silver nitrate titration; weak acid dissociable cyanide by distillation. |

# Potential Hazards

The step-by-step hazard analysis for each task, with the risk ranking, is in the JSEA for that task. The hazards common across this area are:

Contact with sulphuric or hydrochloric acid - burn to skin or eyes

Contact with sodium hydroxide solution - burn to skin or eyes

Hydrogen cyanide gas released at the open sample point

Inhalation of acid fume or reagent vapour at the bench

Cuts from broken glassware, and burns from hot glassware

Manual handling injury from carrying sample buckets, containers and composite drums

Cross-contamination from a container or cutter not cleaned between points

Non-representative sample leading to a wrong process decision

Contact with live electrical equipment at a sample pump, cutter or local control station

Splash of cyanide-bearing slurry or solution to skin or eyes

Contact with hot solution or hot surface at the elution, stripping or regeneration circuits

Contact with an agitator, sample cutter, pump or screen that starts without warning

Hearing damage in the mill, screen, compressor and blower areas

Inhalation of or contact with solid reagent - cyanide briquettes, lime or salt

# Personal Protective Equipment

Standard PPE is worn for all normal operations: hard hat, safety glasses, high-visibility long-sleeved shirt and long trousers, safety boots, and gloves appropriate to the task.

The following is added where the task requires it:

Personal HCN gas monitor, calibrated and in test date

Chemical splash goggles and face shield

Nitrile rubber gloves

Nitrile PVC boots

Chemical resistant suit

Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

Acid-resistant gloves, apron and face shield

Chemical splash goggles

# Procedure

## Introduction

| Warning: | many operations on the gold processing plant are automated and equipment may start or stop without warning. It is the operator's responsibility to ensure the safety of all personnel working on the plant at all times. |
| --- | --- |

## Procedure and Implementation Steps

The documents below are the controlled method for each task. This procedure does not repeat them.

| JSEA | Work instruction | Task | Class |
| --- | --- | --- | --- |
| JSEA-PRO-MET-206-001 | SWI-PRO-MET-206-001 | Met Lab Chemical Receipt, Storage and Decanting | Safety-critical |
| JSEA-PRO-MET-206-002 | SWI-PRO-MET-206-002 | Fume Cupboard Operation, Scrubber Check and Cleaning | Safety-critical |
| JSEA-PRO-MET-206-003 | SWI-PRO-MET-206-003 | Balance Verification and Calibration Check | Operating |
| JSEA-PRO-MET-206-004 | SWI-PRO-MET-206-004 | pH, DO and Conductivity Meter Calibration | Operating |
| JSEA-PRO-MET-206-005 | SWI-PRO-MET-206-005 | Cyanide Solution Handling and Spill Response | Safety-critical |
| JSEA-PRO-MET-206-006 | SWI-PRO-MET-206-006 | Acid Spill Response and Neutralisation | Safety-critical |
| JSEA-PRO-MET-206-007 | SWI-PRO-MET-206-007 | Met Lab Waste, Residue and Effluent Disposal | Safety-critical |
| JSEA-PRO-MET-206-008 | SWI-PRO-MET-206-008 | Glassware Cleaning and Drying | Operating |
| JSEA-PRO-MET-206-009 | SWI-PRO-MET-206-009 | Bottle Roller, Bench Mill and Agitator Operation | Operating |
| JSEA-PRO-MET-206-010 | SWI-PRO-MET-206-010 | Drying Oven and Muffle Furnace Operation | Safety-critical |
| JSEA-PRO-MET-206-011 | SWI-PRO-MET-206-011 | Met Lab Emergency Response - Power Failure and Evacuation | Safety-critical |
| JSEA-PRO-MET-206-012 | - | Met Lab QA/QC - Standards, Duplicates, Blanks and Repeats | Operating |
| JSEA-PRO-MET-206-013 | SWI-PRO-MET-206-013 | Ductless Fume Cabinet Filter Change and Breakthrough Check | Safety-critical |
| JSEA-PRO-MET-206-014 | SWI-PRO-MET-206-014 | Met Lab Equipment Receipt, Installation and Commissioning Handover | Operating |

## Schedule

| Task | Frequency |
| --- | --- |
| 206-001 Met Lab Chemical Receipt, Storage and Decanting | Every delivery, and every decant. |
| 206-002 Fume Cupboard Operation, Scrubber Check and Cleaning | Every use; scrubber checked at the start of every shift. |
| 206-003 Balance Verification and Calibration Check | Daily verification before use; full calibration at the external interval. |
| 206-004 pH, DO and Conductivity Meter Calibration | Every shift before use, and after any electrode change. |
| 206-005 Cyanide Solution Handling and Spill Response | Every cyanide handling task; drill at the set interval. |
| 206-006 Acid Spill Response and Neutralisation | Every acid handling task; drill at the set interval. |
| 206-007 Met Lab Waste, Residue and Effluent Disposal | Every day, and at the end of every test. |
| 206-008 Glassware Cleaning and Drying | After every use. |
| 206-009 Bottle Roller, Bench Mill and Agitator Operation | Every testwork run using them. |
| 206-010 Drying Oven and Muffle Furnace Operation | Every drying and ashing operation. |
| 206-011 Met Lab Emergency Response - Power Failure and Evacuation | On any power failure or evacuation; drill at the set interval. |
| 206-012 Met Lab QA/QC - Standards, Duplicates, Blanks and Repeats | Every analytical batch; charts reviewed weekly and reported monthly. |
| 206-013 Ductless Fume Cabinet Filter Change and Breakthrough Check | Breakthrough check monthly; filter change on indication or at the manufacturer's interval. |
| 206-014 Met Lab Equipment Receipt, Installation and Commissioning Handover | Once, at laboratory establishment; then on every new instrument. |

## Flowchart

Not reproduced here. Each work instruction carries its own step sequence.

# Training

Every person carrying out a task in this area is trained and signed off against the work instruction for that task, and has signed on to the matching JSEA at a communication session.

Site induction, current.

Cyanide awareness and gas detection competency where the area requires it.

Working at height competency where the task requires it.

Competency in the determinations, instruments and calculations used in the task.

Refresher at the interval set by the site training matrix, and after any change to the instruction.

# Document Traceability

The three-digit suffix is shared between the hazard analysis and the work instruction, so SWI-PRO-MET-206-011 is the work instruction for JSEA-PRO-MET-206-011.

| Item | Count |
| --- | --- |
| Tasks in area 206 | 14 |
| Hazard analyses issued | 14 |
| Work instructions issued | 13 |
| Desk-based tasks with a JSEA but no work instruction | 1 |
| Safety-critical | 8 |
| Issued but NOT FOR USE | 0 |

# Review Criteria

This Procedure and the documents beneath it shall be reviewed as follows:

Every 12 months; or

Immediately after any incident, plant change, equipment change or change to the schedule.

# Legislation & References

## Legislation

Queensland Mining and Quarrying Safety and Health Act 1999

Queensland Mining and Quarrying Safety and Health Regulation 2017

## Source Documents

The source documents and notes held for the tasks in this area.

KBK-MIR-LAB-WET-SOP-025 chemical issue and WET-SOP-047 acid transfer

ILS quotation Q0023979 - 160 L flammable, 160 L corrosive and 60 L toxic/poisons cabinets

KBK-MIR-LAB-WET-SOP-054, -059 and ENV-SOP-061 fume cupboard and scrubber

ILS quotation Q0023979 - 2 x 2400 mm Safetyflow fume cupboard rated for perchloric and HF, Model A crossflow scrubber, exhaust fans

KBK-MIR-LAB-PRE-SOP-012 and WET-SOP-013, -014, -015 balance verification

ILS quotation Q0023979 - Class F1 weight set 10 mg to 200 g, GR-200 and BM-22 balances, 150 kg platform balance

KBK-MIR-MP-PRO-MET-SOP-0030 Kalibrasi DO-meter and SOP-0031 Kalibrasi pH-meter, both received 11 Aug 2026

ILS quotation Q0023979 - 2 x pH/ORP meters and 2 x DO meter HI9142

SWI-022 Gas Detector Alarm Response and SWI-025 Cyanide Bearing Spill Response, both issued (plant versions)

KBK-MIR-MP-PRO-OPE-SOP-0076 Fasilitas Cyanide, received 11 Aug 2026

ILS quotation Q0023979 - safety shower with eye wash, 3 x eye wash stations, 120 L hazchem spill kit

KBK-MIR-LAB-WET-SOP-051 Acid spill handling

ILS quotation Q0023979 - 120 L hazchem spill kit, safety shower and eye wash stations

KBK-MIR-LAB-ENV-SOP-058 sludge transfer and ENV-SOP-106 container neutralisation

ILS quotation Q0023979 - 750 L tipping bin, 2 x 240 L Sulo, 60 L drum, crossflow scrubber with pH control

KBK-MIR-LAB-WET-SOP-049 and ENV-SOP-060 glassware drying and washing

ILS quotation Q0023979 - Bottle Roller MkII 12-place with guarding, 4 x MS8 magnetic stirrers

KBK-MIR-MP-PRO-OPE-SOP-0003 Pulverized Bottle Roll Test

KBK-MIR-LAB-PRE-SOP-003 Pengeringan Sampel and WET-SOP-024 ashing furnace

ILS quotation Q0023979 - 2 x drying oven 560 L with 70 trays, ashing furnace 3.5 kW, muffle furnace SF-3-SD

KBK-MIR-LAB-WET-SOP-050 Evacuation on sudden power failure

ILS quotation Q0023979 - 2 x Riello UPS 2200VA

No source held. The KBK set carries individual methods but no QA/QC system.

ILS quotation Q0023979 - no certified reference materials included, and the quote states plainly that no standards or consumables come with the Eltra CS-i

ILS quotation Q0023979 - Purair 30-PP ductless cabinet with carbon filters

No procedure was held for this cabinet before this revision

ILS quotation Q0023979 - the asset list, ex works, 12 to 18 week lead time, with 10 per cent of value payable against commissioning

The quotation's stated validity expired 16 February 2026 - it is a quotation, not a purchase order