Safe Work Instruction

WAD Cyanide - Distillation and Determination

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Nitrile PVC boots

- Chemical resistant suit, acid-resistant gloves and apron

- Heat-resistant gloves for hot glassware

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-LAB-304-008 - WAD Cyanide - Distillation and Determination |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-LAB-304-008 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Determine weak acid dissociable (WAD) cyanide by distillation with an acetate buffer at pH 4.5. The HCN released is carried by air into a sodium hydroxide absorber, which is then analysed. WAD cyanide counts free cyanide plus the cyanide held in weak metal complexes such as zinc, copper and nickel, which a free cyanide test does not see. Frequency: To be set once a distillation train is purchased. This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO215 - Laboratory Environmental and Cyanide Analysis. Read that procedure before carrying out this task for the first time. Register line: SWI-PRO-LAB-304-008 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-LAB. The matching hazard analysis is JSEA-PRO-LAB-304-008. Status: No cyanide distillation train is included in ILS quote Q0023979. WAD cyanide is a licence-condition analysis for a cyanide plant - resolve before commissioning. This instruction is not for use until the train is purchased and the items marked (TBC) are checked against the method. This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Hydrogen cyanide gas generated in the distillation flask by the acid buffer — All work in the fume cupboard. Train leak tested before use. Air drawing through the absorber before any buffer goes in. Personal HCN monitor worn. Fixed detection healthy - alarm 5 ppm, high-high 10 ppm. Antidote kit on site and a trained first aider on shift. Contact with the acid buffer, zinc acetate or sodium hydroxide — Face shield, goggles, acid-resistant gloves and apron. Acid added to water, never the reverse. Every volume drawn with a pipette filler. Suck-back or pressure build-up in the train — Absorber watched for steady bubbling throughout. Heat off at the first sign of suck-back. Air kept drawing. A closed system is never heated. Burns from the boiling flask and heating mantle, and cuts from broken glassware — Flask clamped and handled hot only with heat-resistant gloves. Train cooled before it is dismantled. Glassware inspected for star cracks and discarded if damaged. Wrong result from a leak, an incomplete distillation or carry-over — Leak test before every batch. Distillation run for its full time. A blank and a check standard run with every batch. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-LAB-304-008 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - HCN detection not healthy - Train not leak tested - Cooling water not flowing |  | - Confirm the fume cupboard airflow, the scrubber and its pH dosing are running, and the sash is at the marked working height. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the Shift Supervisor knows the task is starting and a second person is in the area. - Confirm the personal HCN monitor is calibrated, in test date and switched on, and the fixed detection is healthy. - Confirm the cyanide antidote kit is on site and a trained first aider is on shift. - Confirm the train has been leak tested, dried and cleaned to SWI-PRO-LAB-304-010. - Confirm the reagents were made up to SWI-PRO-LAB-304-019 and are in date. - Inspect the flask, condenser and absorber for star cracks and discard anything damaged. Inspect the heating mantle lead and controller. Equipment Required : - Cyanide distillation train – TBC (not yet purchased) - Heating mantle with controller - Air or vacuum source with flow control – TBC - Gas absorber with sodium hydroxide absorbing solution – strength TBC - Acetate buffer pH 4.5 and zinc acetate – to SWI-PRO-LAB-304-019 - Pipettes and pipette filler - Volumetric flasks - Cooling water supply - Heat-resistant gloves - Fume cupboard |  |  |  |  |  |  |  |  |  |  |  |  |
| - Apparatus Set-up - Leak in the train — HCN to the cupboard - Absorber not drawing — HCN has nowhere to go - Condenser dry — HCN and steam carried past |  | - Assemble the train in the fume cupboard and clamp the flask, condenser and absorber. - Fill the absorber with the sodium hydroxide absorbing solution (volume TBC). - Turn on the cooling water and confirm it is flowing through the condenser before any heat is applied. - Start the air draw and confirm steady bubbling through the absorber. If it does not bubble, the train leaks — stop and find the leak. CAUTION: Air must be drawing through the absorber before the buffer goes in. Otherwise the HCN has nowhere to go but out. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Sample Charge and Distillation - HCN release when the buffer is added - Splash of acid buffer or sample - Suck-back from the absorber |  | - Pipette the sample (volume TBC) into the distillation flask. - With the air drawing, add the zinc acetate and the acetate buffer through the reagent inlet. Never lift the flask stopper once the buffer is in. - Rinse the inlet with a little deionised water. - Switch on the heating mantle and bring the flask to a steady reflux. - Distil for the full time the method states (TBC) with the air drawing steadily. - Stay with the train while it is heating. Watch the absorber for steady bubbling. - Run a blank and a check standard through the same sequence with every batch. CAUTION: If the absorber stops bubbling or liquid starts to suck back, switch off the heat, keep the air drawing and call the Shift Supervisor. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Cool Down and Determination - HCN left in the flask when the air is stopped early - Burns from the hot flask and mantle - Calculation error |  | - Switch off the heat and keep the air drawing while the flask cools (time TBC). This sweeps the last HCN into the absorber. - Stop the air only when the flask is cool. - Transfer the absorber solution to a volumetric flask and make up to the mark with deionised water. - Determine cyanide in the absorber solution by titration to SWI-PRO-LAB-304-005, or by colorimetric reading to SWI-PRO-LAB-304-006 once it is in use. - Calculate WAD cyanide in the sample from the absorber result, the absorber volume and the sample volume (TBC). - Have the calculation cross-checked before the result is reported. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Shutdown and Clean Up - Acidic flask residue still releasing HCN - Cyanide-bearing solution to the general drain - Results not traceable |  | - Make the flask residue alkaline with sodium hydroxide before dismantling. It is acidic and still holds cyanide. - Empty the flask residue and spent absorber solution to cyanide waste. - Leave the mantle to cool before moving it. Rinse and clean the train to SWI-PRO-LAB-304-010. - Wash down the bench and the fume cupboard working surface. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the train identity. Any step not completed is recorded with the reason. - Any control in Part 2 found not in place is reported to the Shift Supervisor and recorded. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| HM-PRC-VXX-PRO215 | Laboratory Environmental and Cyanide Analysis |
| JSEA-PRO-LAB-304-008 | JSEA — WAD Cyanide - Distillation and Determination |
| HM-MMM-WHS-REG-XXX-R00-SWI-SWMS | SWI / SWMS Register — tab SWI-LAB |
| SWI-PRO-LAB-304-005 | Free Cyanide by Titration - Silver Nitrate and Rhodanine |
| SWI-PRO-LAB-304-006 | Free Cyanide by Direct Colorimetric Reading |
| SWI-PRO-LAB-304-010 | Distillation Apparatus Leak Test, Drying and Cleaning |
| SWI-PRO-LAB-304-019 | Preparation of Cyanide Analysis Reagents - Pre and Post Distillation |
| ILS Quote Q0023979 | Laboratory equipment quote — no cyanide distillation train included |
| TBC | WAD cyanide distillation method and train operating manual (not yet purchased) |

In the event of an emergency

1. HCN Gas Alarm If the personal monitor or fixed detection alarms (5 ppm alarm, 10 ppm high-high), stop work, leave the area upwind and call CH19 “EMERGENCY”. Do not re-enter until the Shift Supervisor clears the area. 2. Leak or Suck-back in the Train Switch off the heat at the controller, keep the air drawing, lower the sash and step back. Call the Shift Supervisor. Do not reach in until the HCN monitor reads clear. 3. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 4. Skin or Eye Contact Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. 5. Burns Cool the burn under cool running water for 20 minutes. Do not apply ice. Seek first aid.