Safe Work Instruction

Met Lab Sample Receipt, Logging and Storage

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, nitrile rubber gloves and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-204-001 - Met Lab Sample Receipt, Logging and Storage |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-204-001 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Receive samples into the metallurgical laboratory with a chain of custody that holds. A sample that cannot be traced back to a point, a time and a shift is worthless no matter how well it is analysed. Area: 204 Met Lab - Physical Testwork. Points and equipment covered: Met lab bench - physical testwork bench and instrument Frequency: Every delivery from the plant. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. Hydrogen cyanide gas released at the open sample point — Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Manual handling injury from carrying sample buckets, containers and composite drums — Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. Non-representative sample leading to a wrong process decision — Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-204-001 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-204-001 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - Sample register or LIMS entry - Balance - Storage racks and bins - Labels and marker - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Delivery Check - Sample missing or not listed - Leaking or unlabelled sample accepted |  | - Check the delivery against the handover sheet - every sample listed is present and every sample present is listed. - Reject and record any sample that is unlabelled, illegible, leaking or clearly the wrong material. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Logging and Numbering - Sample not traceable to point, time and shift - Receipt mass not recorded |  | - Log each sample against its point, date, time, shift and the analyses requested. - Assign a laboratory number and mark it on the container. - Weigh the sample where the method requires a receipt mass. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Storage and Handover - High-grade sample stored with low-grade - Manual handling of heavy containers - Handover not signed |  | - Store to the storage rack for its stream. Keep high-grade streams separate from low-grade. - Sign the handover and return a copy to the plant. CAUTION: Two-person lift or a trolley for containers above the site manual handling limit. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| JSEA-PRO-MET-204-001 | JSEA — Met Lab Sample Receipt, Logging and Storage |
| KBK-MIR-LAB-PRE-SOP-001 and PRE-SOP-002 | Penerimaan Sampel and Sortir Sampel |
| 4034-PR-PRO-002 | Sampling Protocol |

In the event of an emergency

1. HCN Gas Alarm If the personal monitor or fixed detection alarms (5 ppm alarm, 10 ppm high-high), stop work, leave the area upwind and call CH19 “EMERGENCY”. Do not re-enter until the Shift Supervisor clears the area. 2. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 3. Skin or Eye Contact with Slurry Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 4. Slip, Fall or Lifting Injury Call CH19 “EMERGENCY”, call 000 if the injury is serious, and notify the Shift Supervisor.