Safe Work Instruction

Centrifuge Operation for Solution Clarification

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, nitrile rubber gloves and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-204-019 - Centrifuge Operation for Solution Clarification |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-204-019 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Clarify a solution sample by centrifuge where filtration is impractical or too slow. Used where a fine or gelatinous solid blinds a filter paper. Area: 204 Met Lab - Physical Testwork. Points and equipment covered: Met lab bench - physical testwork bench and instrument Frequency: As required by the analysis schedule. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Contact with an agitator, sample cutter, pump or screen that starts without warning — Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. Cuts from broken glassware, and burns from hot glassware — Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. Splash of cyanide-bearing slurry or solution to skin or eyes — Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. Hydrogen cyanide gas released at the open sample point — Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. Non-representative sample leading to a wrong process decision — Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-204-019 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-204-019 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - Ohaus FC5706 centrifuge - 15 mL centrifuge tubes - Balance for tube pairing - Timer - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Tube Preparation and Loading - Cracked tube failing at speed - Unbalanced rotor - Splash of cyanide-bearing solution |  | - Inspect the tubes for cracks before loading. A cracked tube fails at speed. - Fill opposing tubes to the same level and balance them on the balance to within the manufacturer's tolerance. An unbalanced rotor is the main failure mode of a centrifuge. - Load opposing positions symmetrically. Never run with a single tube. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Running the Centrifuge - Lid opened while spinning - Rotor braked by hand |  | - Close and latch the lid before starting. - Run at the speed and for the period the method specifies. - Allow the rotor to come to a complete stop on its own. Never brake it by hand. CAUTION: Never open the lid or touch the rotor until it has stopped on its own. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Decanting and Records - Pellet disturbed - Run conditions not recorded |  | - Decant the clarified supernatant without disturbing the pellet. - Record speed, time and the tube set used. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| JSEA-PRO-MET-204-019 | JSEA — Centrifuge Operation for Solution Clarification |
| ILS quotation Q0023979 | Ohaus FC5706 centrifuge and 15 mL tubes |
| Note | No procedure was held for this instrument before this revision |

In the event of an emergency

1. HCN Gas Alarm If the personal monitor or fixed detection alarms (5 ppm alarm, 10 ppm high-high), stop work, leave the area upwind and call CH19 “EMERGENCY”. Do not re-enter until the Shift Supervisor clears the area. 2. Contact with Moving Equipment If a person is caught by or struck by the centrifuge, hit the nearest emergency stop, call CH19 “EMERGENCY” and call 000. Do not approach the equipment until it is isolated and locked out. 3. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 4. Skin or Eye Contact with Slurry Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 5. Cuts from Broken Glassware Apply pressure to stop the bleeding, get a trained first aider, and report to the Shift Supervisor. Call CH19 “EMERGENCY” and 000 for a serious cut.