Safe Work Instruction

Preparation of Standard and Working Solutions

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, acid-resistant and nitrile rubber gloves, and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-205-017 - Preparation of Standard and Working Solutions |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-205-017 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Make up the standard and working solutions the laboratory runs on. Every titration result is only as good as the standard it was titrated against. An expired phenolphthalein indicator or silica gel may be re-qualified by the Martabe test. No other reagent is re-qualified. Area: 205 Met Lab - Chemical and Leach Testwork. Points and equipment covered: Met lab bench - physical testwork bench and instrument Frequency: As the solution register requires; standards re-standardised at the set interval. Draft for review: Steps 5 and 6 follow the Martabe work instruction Testwork QC Expired Reagent (28/03/2025). Re-qualification is limited to phenolphthalein indicator and silica gel, never standards, primary reagents or cyanide, and needs the approval of the Metallurgy Superintendent. JSEA-PRO-MET-205-017 must be updated for the added hazards before this instruction is approved. Open items: (1) Metallurgy Superintendent approval of the re-qualification policy; (2) the silica gel drying temperature is not stated - take it from the safety data sheet (typically about 110 to 120 deg C); (3) the Martabe Reagent Lifetime Management System (DOC-IV-MET-CHH-SOP-00039) has not been received. This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Contact with sulphuric or hydrochloric acid - burn to skin or eyes — Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Contact with sodium hydroxide solution - burn to skin or eyes — Full chemical PPE including face shield and chemical-resistant gloves. Sample drawn slowly from the designated valve. Safety shower and eyewash within reach and proven flowing. Hydrogen cyanide gas released at the open sample point — Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Inhalation of acid fume or reagent vapour at the bench — Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Cuts from broken glassware, and burns from hot glassware — Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. Non-representative sample leading to a wrong process decision — Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. [JSEA to confirm] 96 per cent ethanol (flammable) near the drying oven, and dilute acid and alkali used for indicator checks — Ethanol used in small volumes in the fume cupboard, away from the oven and ignition sources, container closed after use. Dilute acid and alkali handled with splash goggles and chemical-resistant gloves. Waste to the correct stream. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-205-017 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-205-017 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm a second person is in the laboratory for any cyanide work. Cyanide work is not done alone. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - Analytical balance 0.1 mg and Class F1 weight set - Volumetric flasks and pipettes - Distilled or RO/DI water - Fume cupboard - Reagent chemicals in date - Amber bottles and labels - Drying oven - 96 per cent ethanol, diluted NaOH, diluted HCl or H2SO4 - Phenolphthalein powder and silica gel for re-qualification - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Weighing the Standard - Balance not verified - Reagent of unknown age - Cyanide or silver nitrate contact |  | - Check the balance verification is current before weighing any standard. - Use only reagents in date and correctly labelled. A reagent of unknown age does not make a standard. The only exception is the re-qualification of phenolphthalein and silica gel in Steps 5 and 6. - Weigh the primary standard on the analytical balance and record the exact mass. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Making Up and Standardising - Volume made up off temperature - Factor not recorded |  | - Dissolve and make up to the mark in a volumetric flask at the reference temperature. - Standardise the solution against a primary standard where the method requires it, and record the factor. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Labelling, Storage and Register - Unlabelled solution used - Silver nitrate exposed to light |  | - Label every solution with its identity, concentration, preparation date, expiry and the preparer's initials. An unlabelled solution is discarded, not guessed at. - Store light-sensitive solutions such as silver nitrate in amber bottles away from light. - Record the preparation in the solution register and re-standardise at the set interval. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Re-qualifying Expired Phenolphthalein - Flammable ethanol near the oven - Dilute acid or alkali splash |  | - Inspect it visually. It stays usable only if it is not cloudy and has no deposit. - Dissolve a little phenolphthalein powder in 96 per cent ethanol in the fume cupboard. Usable phenolphthalein dissolves completely. - Make up the indicator solution as SOP-0050. - Add 2 to 3 drops to diluted NaOH - usable only if the solution turns pink. Add 2 to 3 drops to diluted HCl or H2SO4 - usable only if the solution stays clear. - If any test fails, discard the phenolphthalein. If all pass, record the test date and the new expiry date, 1 year after the test. CAUTION: Ethanol is flammable. Small volumes in the fume cupboard, away from the oven and any ignition source. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Re-qualifying Expired Silica Gel - Hot items from the oven |  | - Inspect it visually. If the colour has moved away from orange, it has absorbed moisture and its effectiveness has dropped. - Dry the silica gel in the oven for 2 hours at the temperature on the safety data sheet (see Open items). - After drying, check whether the colour returns to orange. If it does not, discard the silica gel. Record the result. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| JSEA-PRO-MET-205-017 | JSEA — Preparation of Standard and Working Solutions |
| KBK-MIR-LAB-WET-SOP-043, -044 and -045 | Standard solutions |
| KBK-MIR-MP-PRO-MET-SOP-0043, -0050, -0051, -0052 and -0053 | AgNO3 0.01 M, phenolphthalein, rhodanine, 10 per cent sulfamic acid and KCl for the cyanide analyser (received 11 Aug 2026) |
| ILS quotation Q0023979 | Class F1 weight set 10 mg to 200 g and GR-200 0.1 mg balance |
| Martabe WI (28/03/2025) | Testwork QC Expired Reagent - source of Steps 5 and 6 |
| DOC-IV-MET-CHH-SOP-00039 | Martabe Reagent Lifetime Management System - not received |

In the event of an emergency

1. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 2. Skin or Eye Contact with Reagent Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 3. HCl Gas or Acid Contact Leave the area upwind. Flush skin or eyes with copious water at the safety shower or eyewash for a minimum of 15 minutes. If breathing is affected call CH19 “EMERGENCY” and call 000. 4. Burns Cool the burn under cool running water for 20 minutes. Do not apply ice. Seek first aid. 5. Cuts from Broken Glassware Apply pressure to stop the bleeding, get a trained first aider, and report to the Shift Supervisor. Call CH19 “EMERGENCY” and 000 for a serious cut.