Safe Work Instruction

pH, DO and Conductivity Meter Calibration

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, acid-resistant and nitrile rubber gloves, and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-206-004 - pH, DO and Conductivity Meter Calibration |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-206-004 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Calibrate the laboratory and field pH, dissolved oxygen and conductivity meters. pH governs cyanide safety as well as process control. An uncalibrated pH meter is a safety issue, not just a data issue. Area: 206 Met Lab - Facility, Safety and Quality. Points and equipment covered: Met lab bench - physical testwork bench and instrument Plant tanks - dissolved oxygen measured in slurry Frequency: Every shift before use, and after any electrode change. Draft for review: Steps 3 and 4 follow the Martabe work instructions pH Meter TPS Cube Calibration (DOC-3-MET-PMC-WIN-00087-IE) and Operating pH Cube Meter (DOC-3-MET-PMC-WIN-00089-IE), and Steps 5 and 6 follow Operate DO Meter Portable (DOC-3-MET-PRS-WIN-00055-IE), v1.0, 25/12/2024. The TPS pH cube meter is adopted as at Martabe. pH calibration stays at three points (pH 4, 7 and 10) with the slope and offset recorded; the Martabe two-point calibration is not used. The Martabe DO method is applied to the Hanna HI9142 meter without the Oxyguard menu steps. JSEA-PRO-MET-206-004 must be updated for the added hazards before this instruction is approved. Open items: (1) confirm the TPS pH cube meter is on site, or procure it, and add it to the asset list; (2) take the DO calibration acceptance limit in air from the HI9142 manual - the Martabe value of 6 to 8 has no unit and is not used; (3) confirm from the TPS cube manual which control sets pH 4 - the source says "the same as pH 7"; (4) set the acceptable pH electrode slope range; (5) confirm the Mt. Morgan records for DO results and pH calibration - the Martabe Daily Task Checklist (DOC-FRM-00052-EN) and form DOC-4-MET-PMC-DFR-00142-EN are source names; (6) conductivity standards are not held - conductivity meters cannot be calibrated until they are. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Contact with sulphuric or hydrochloric acid - burn to skin or eyes — Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Contact with sodium hydroxide solution - burn to skin or eyes — Full chemical PPE including face shield and chemical-resistant gloves. Sample drawn slowly from the designated valve. Safety shower and eyewash within reach and proven flowing. Cuts from broken glassware, and burns from hot glassware — Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. Contact with live electrical equipment at a sample pump, cutter or local control station — No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. Non-representative sample leading to a wrong process decision — Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. [JSEA to confirm] Contact with the sample and splash to the eyes at the pH meter — Chemical-resistant gloves, glasses and safety boots. Eyewash proven flowing; flush the eyes and skin under running water at once. [JSEA to confirm] Buffer splash and spill, and a scratch from the screwdriver at the pH meter — Chemical-resistant gloves, glasses and safety boots; pour slowly; spill kit at hand and clean spills at once; eyewash proven flowing. Check the screwdriver is in good condition and watch hand position when turning the slope control; clean and treat any scratch. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-206-004 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-206-004 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Take 5 and prepare the meters, the buffer solutions, a beaker and a small bucket. Check the equipment is in good condition and the buffers are in date. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - pH buffers 4, 7 and 10, in date - TPS pH cube meter with pH and temperature sensors - pH/ORP meters (ILS quotation) - Hanna HI9142 DO meters with their manual and a spare membrane - Conductivity standards - NOT HELD - Flat screwdriver for the calibrate and slope controls - Beaker and small bucket - Distilled water, soft tissue and electrode storage solution - Meter log and DO checklist - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - pH Meter Calibration - Three Point - Out-of-date buffer - Buffer splash to the eyes - Electrode slope out of range |  | - Check the buffer expiry date. An out-of-date buffer calibrates the meter to the wrong value. - Rinse the electrode with distilled water and blot dry between buffers. - Calibrate the pH meter across pH 4, 7 and 10 and record the slope and offset. Two-point calibration is not used. - Reject the electrode if the slope is outside the acceptable range. - Record the temperature at calibration - pH is temperature dependent. |  |  |  |  |  |  |  |  |  |  |  |  |
| - TPS Cube pH Calibration - Buffer splash and spill - Scratch from the screwdriver - Calibration not recorded |  | - pH 7: wash the pH probe with water. Pour pH 7 buffer into the beaker and put the pH and temperature sensors in it. Wait until the reading is steady at 7. Adjust to 7 with the 'calibrate' knob. Clean the sensors in water. - pH 10: pour pH 10 buffer into the beaker and put the sensors in. Wait until the reading is steady. Adjust to 10 with the slope control, using the screwdriver. Clean the sensors in water. - pH 4: pour pH 4 buffer into the beaker and put the sensors in. Wait until the reading is steady and adjust to 4 as the meter manual directs (see Open items). Clean the sensors in water. - Record the readings, the slope and offset, and the condition of the meter before and after calibration in the pH calibration logsheet. CAUTION: Pour buffers slowly and clean any spill at once. Check the screwdriver is in good condition and watch hand position on the slope control. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Measuring pH with the TPS Cube - Sample contact and splash to the eyes - Probe allowed to dry out - Cyanide-bearing sample to the wrong waste stream |  | - Calibrate the meter first (Step 3). - Set the pH cube to pH mode. Plug the pH sensor and the temperature sensor into their sockets. Remove the cap from the pH probe. - Rinse the probe with clean water and dry it. - Dip both sensors in the solution to be measured. Wait until the reading is stable, then record it. - Rinse the probe with clean water and keep it in water so it stays wet. - Pour the sample to the site waste stream for its classification. Cyanide-bearing samples go to the cyanide waste stream, never the general drain. |  |  |  |  |  |  |  |  |  |  |  |  |
| - DO Meter Calibration in Air - HI9142 - Damaged or dry membrane - Reading outside the manual limit - Meter used out of calibration |  | - Switch on the DO meter and calibrate it to 100 per cent air saturation, in air or against air-saturated water, following the HI9142 manual. If the membrane needs wiping or changing, do it first. - Check the reading in air against the acceptance limit in the HI9142 manual. Outside the limit, calibrate again; if it is still outside, change the membrane and electrolyte or take the meter out of service. - Record the temperature at calibration - DO is temperature dependent. |  |  |  |  |  |  |  |  |  |  |  |  |
| - DO Measurement in Slurry - HCN at an open cyanide tank - Slurry splash when lifting the probe - Trip on the meter cable |  | - At a cyanide-bearing tank apply the cyanide controls of SWI-PRO-MET-201-003: personal HCN monitor on, approach from upwind, second person in sight. - Tidy the DO meter cable and do not hurry when walking. - Put the DO probe in the tank and wait until the reading is stable. Record the reading, then lift the probe out of the slurry slowly. - Rinse the probe with clean water. - Enter the DO result in the DO checklist (see Open items). CAUTION: Cyanide-bearing slurry. Personal HCN monitor on, upwind approach and a second person in sight. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Electrode Storage and Meter Log - Electrode stored in distilled water - Calibration not traceable to the meter |  | - Store electrodes in the correct storage solution, never in distilled water. - Record every calibration in the meter log against the meter and electrode identity. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| JSEA-PRO-MET-206-004 | JSEA — pH, DO and Conductivity Meter Calibration |
| KBK-MIR-MP-PRO-MET-SOP-0030 | Kalibrasi DO-meter (received 11 Aug 2026) |
| KBK-MIR-MP-PRO-MET-SOP-0031 | Kalibrasi pH-meter (received 11 Aug 2026) |
| ILS quotation Q0023979 | 2 x pH/ORP meters and 2 x DO meter HI9142 |
| DOC-3-MET-PMC-WIN-00087-IE v1.0 | Martabe WI pH Meter TPS Cube Calibration (25/12/2024) - source of Step 3 |
| DOC-3-MET-PMC-WIN-00089-IE v1.0 | Martabe WI Operating pH Cube Meter (25/12/2024) - source of Step 4 |
| DOC-3-MET-PRS-WIN-00055-IE v1.0 | Martabe WI Operate DO Meter Portable (25/12/2024) - source of Steps 5 and 6 |
| DOC-2-MET-PRM-SOP-00037-IE, DOC-2-MET-MEL-SOP-00087-IE, DOC-IV-MET-CHH-SOP-00039 | Martabe SOPs referenced by the Martabe WIs - not received |
| SWI-PRO-MET-201-003 | Leach Tank Sampling Round - cyanide controls at the tank |

In the event of an emergency

1. HCN Gas Alarm If the personal monitor or fixed detection alarms (5 ppm alarm, 10 ppm high-high), stop work, leave the area upwind and call CH19 “EMERGENCY”. Do not re-enter until the Shift Supervisor clears the area. 2. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 3. Skin or Eye Contact with Reagent Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 4. HCl Gas or Acid Contact Leave the area upwind. Flush skin or eyes with copious water at the safety shower or eyewash for a minimum of 15 minutes. If breathing is affected call CH19 “EMERGENCY” and call 000. 5. Cuts from Broken Glassware Apply pressure to stop the bleeding, get a trained first aider, and report to the Shift Supervisor. Call CH19 “EMERGENCY” and 000 for a serious cut.