Safe Work Instruction

Bottle Roller, Bench Mill and Agitator Operation

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet, safety glasses and hearing protection

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, acid-resistant and nitrile rubber gloves, and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-206-009 - Bottle Roller, Bench Mill and Agitator Operation |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-206-009 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Operate the bottle roller, bench mill and magnetic stirrers safely. These are the rotating machines in the laboratory and the guarding is what keeps hands out of them. Area: 206 Met Lab - Facility, Safety and Quality. Points and equipment covered: Met lab bench - physical testwork bench and instrument Frequency: Every testwork run using them. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Contact with an agitator, sample cutter, pump or screen that starts without warning — Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. Hydrogen cyanide gas released at the open sample point — Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Splash of cyanide-bearing slurry or solution to skin or eyes — Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. Hearing damage in the mill, screen, compressor and blower areas — Hearing protection mandatory in signed areas. Double protection at the mill and compressor. Time at the point kept to what the task needs. Cuts from broken glassware, and burns from hot glassware — Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. Manual handling injury from carrying sample buckets, containers and composite drums — Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-206-009 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-206-009 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - Bottle Roller MkII 12-place with guarding - 4 x MS8 magnetic stirrers - Bench mill - Bottles, caps and seals - Timer - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Bottle Roller - Hands caught in the rollers - Bottle opening and spraying cyanide slurry - Unbalanced load |  | - Confirm the guard is in place and the guard interlock stops the machine before loading anything. - Check every bottle cap and seal before loading. A bottle that opens on the roller sprays cyanide slurry across the room. - Load bottles evenly across the rollers so the load is balanced. - Set the speed and the timer to the method and start. - Never reach in to adjust a bottle while the rollers are turning. Stop the machine first. CAUTION: Stop the machine before reaching in. The guard interlock must stop the rollers. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Bench Mill - Contact with the rotating mill - Hearing damage at the mill - Lid opened while turning |  | - For the bench mill, confirm the charge and the lid clamp before starting, and do not open it until it has stopped. CAUTION: Wear hearing protection at the bench mill. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Magnetic Stirrers and Clean-up - Liquid thrown by the vortex - Heated stirrer left unattended |  | - For magnetic stirrers, set the speed so the vortex does not throw liquid, and never leave a heated stirrer unattended. - Clean the rollers and the bench after every run. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| JSEA-PRO-MET-206-009 | JSEA — Bottle Roller, Bench Mill and Agitator Operation |
| ILS quotation Q0023979 | Bottle Roller MkII 12-place with guarding, 4 x MS8 magnetic stirrers |
| KBK-MIR-MP-PRO-OPE-SOP-0003 | Pulverized Bottle Roll Test |

In the event of an emergency

1. HCN Gas Alarm If the personal monitor or fixed detection alarms (5 ppm alarm, 10 ppm high-high), stop work, leave the area upwind and call CH19 “EMERGENCY”. Do not re-enter until the Shift Supervisor clears the area. 2. Contact with Moving Equipment If a person is caught by or struck by the bottle roller, bench mill or stirrer, hit the nearest emergency stop, call CH19 “EMERGENCY” and call 000. Do not approach the equipment until it is isolated and locked out. 3. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 4. Skin or Eye Contact with Slurry Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 5. Cuts from Broken Glassware Apply pressure to stop the bleeding, get a trained first aider, and report to the Shift Supervisor. Call CH19 “EMERGENCY” and 000 for a serious cut. 6. Slip, Fall or Lifting Injury Call CH19 “EMERGENCY”, call 000 if the injury is serious, and notify the Shift Supervisor.