## Register 206
Kode | No. dokumen | Judul | Area | Klasifikasi | JSEA ID | Rev | Tanggal terbit | Review berikutnya | Document owner | Frekuensi | Jumlah step | Halaman | Padanan Martabe | Status isi | Format Heritage | Nama file
206-001 | SWI-PRO-MET-206-001 | Met Lab Chemical Receipt, Storage and Decanting | 206 Met Lab - Facility, Safety and Quality | Safety-critical | JSEA-PRO-MET-206-001 | A | 11/08/2026 | 11/08/2027 | Processing | Every delivery, and every decant. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-001-Met-Lab-Chemical-Receipt-Storage-and-Decanting-RevA.docx
206-002 | SWI-PRO-MET-206-002 | Fume Cupboard Operation, Scrubber Check and Cleaning | 206 Met Lab - Facility, Safety and Quality | Safety-critical | JSEA-PRO-MET-206-002 | A | 11/08/2026 | 11/08/2027 | Processing | Every use; scrubber checked at the start of every shift; face velocity measured monthly and after any repair to the extraction fan, ductwork or scrubber. | 6 | 7 | Operating Anemometer GM8902 - Step 4-5, ukur di tinggi sash kerja, rata-rata dan bacaan terendah dari 20 bacaan | DRAFT FOR REVIEW (tambahan Martabe) | Sudah | SWI-PRO-MET-206-002-Fume-Cupboard-Operation-Scrubber-Check-and-Cleaning-RevA.docx
206-003 | SWI-PRO-MET-206-003 | Balance Verification and Calibration Check | 206 Met Lab - Facility, Safety and Quality | Operating | JSEA-PRO-MET-206-003 | A | 11/08/2026 | 11/08/2027 | Processing | Daily verification before use; full calibration at the external interval. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-003-Balance-Verification-and-Calibration-Check-RevA.docx
206-004 | SWI-PRO-MET-206-004 | pH, DO and Conductivity Meter Calibration | 206 Met Lab - Facility, Safety and Quality | Operating | JSEA-PRO-MET-206-004 | A | 11/08/2026 | 11/08/2027 | Processing | Every shift before use, and after any electrode change. | 8 | 8 | TPS cube pH meter (kalibrasi tetap 3 titik) - Step 3-4; Operate DO Meter Portable untuk HI9142 tanpa menu Oxyguard - Step 5-6 | DRAFT FOR REVIEW (tambahan Martabe) | Sudah | SWI-PRO-MET-206-004-pH-DO-and-Conductivity-Meter-Calibration-RevA.docx
206-005 | SWI-PRO-MET-206-005 | Cyanide Solution Handling and Spill Response | 206 Met Lab - Facility, Safety and Quality | Safety-critical | JSEA-PRO-MET-206-005 | A | 11/08/2026 | 11/08/2027 | Processing | Every cyanide handling task; drill at the set interval. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-005-Cyanide-Solution-Handling-and-Spill-Response-RevA.docx
206-006 | SWI-PRO-MET-206-006 | Acid Spill Response and Neutralisation | 206 Met Lab - Facility, Safety and Quality | Safety-critical | JSEA-PRO-MET-206-006 | A | 11/08/2026 | 11/08/2027 | Processing | Every acid handling task; drill at the set interval. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-006-Acid-Spill-Response-and-Neutralisation-RevA.docx
206-007 | SWI-PRO-MET-206-007 | Met Lab Waste, Residue and Effluent Disposal | 206 Met Lab - Facility, Safety and Quality | Safety-critical | JSEA-PRO-MET-206-007 | A | 11/08/2026 | 11/08/2027 | Processing | Every day, and at the end of every test. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-007-Met-Lab-Waste-Residue-and-Effluent-Disposal-RevA.docx
206-008 | SWI-PRO-MET-206-008 | Glassware Cleaning and Drying | 206 Met Lab - Facility, Safety and Quality | Operating | JSEA-PRO-MET-206-008 | A | 11/08/2026 | 11/08/2027 | Processing | After every use. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-008-Glassware-Cleaning-and-Drying-RevA.docx
206-009 | SWI-PRO-MET-206-009 | Bottle Roller, Bench Mill and Agitator Operation | 206 Met Lab - Facility, Safety and Quality | Operating | JSEA-PRO-MET-206-009 | A | 11/08/2026 | 11/08/2027 | Processing | Every testwork run using them. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-009-Bottle-Roller-Bench-Mill-and-Agitator-Operation-RevA.docx
206-010 | SWI-PRO-MET-206-010 | Drying Oven and Muffle Furnace Operation | 206 Met Lab - Facility, Safety and Quality | Safety-critical | JSEA-PRO-MET-206-010 | A | 11/08/2026 | 11/08/2027 | Processing | Every drying and ashing operation. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-010-Drying-Oven-and-Muffle-Furnace-Operation-RevA.docx
206-011 | SWI-PRO-MET-206-011 | Met Lab Emergency Response - Power Failure and Evacuation | 206 Met Lab - Facility, Safety and Quality | Safety-critical | JSEA-PRO-MET-206-011 | A | 11/08/2026 | 11/08/2027 | Processing | On any power failure or evacuation; drill at the set interval. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-011-Met-Lab-Emergency-Response---Power-Failure-and-Evacuation-RevA.docx
206-013 | SWI-PRO-MET-206-013 | Ductless Fume Cabinet Filter Change and Breakthrough Check | 206 Met Lab - Facility, Safety and Quality | Safety-critical | JSEA-PRO-MET-206-013 | A | 11/08/2026 | 11/08/2027 | Processing | Breakthrough check monthly; filter change on indication or at the manufacturer's interval. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-013-Ductless-Fume-Cabinet-Filter-Change-and-Breakthrough-Check-RevA.docx
206-014 | SWI-PRO-MET-206-014 | Met Lab Equipment Receipt, Installation and Commissioning Handover | 206 Met Lab - Facility, Safety and Quality | Operating | JSEA-PRO-MET-206-014 | A | 11/08/2026 | 11/08/2027 | Processing | Once, at laboratory establishment; then on every new instrument. | 5 | 6 | Tidak ada padanan | Final (isi tanpa tambahan) | Sudah | SWI-PRO-MET-206-014-Met-Lab-Equipment-Receipt-Installation-and-Commissioning-Handover-RevA.docx

## Draft dan open items
No. dokumen | Judul | Jenis | Isi catatan
SWI-PRO-MET-206-002 | Fume Cupboard Operation, Scrubber Check and Cleaning | Draft | Draft for review: Steps 4 and 5 follow the Martabe work instruction Operating Anemometer GM8902 (DOC-3-MET-PMC-WIN-00130-IE, v1.0, 25/12/2024). The face velocity is measured at the marked working sash height, and the average and the lowest of 20 readings are recorded. JSEA-PRO-MET-206-002 must be updated for the added hazards before this instruction is approved.
SWI-PRO-MET-206-002 | Fume Cupboard Operation, Scrubber Check and Cleaning | Open items | Open items: (1) take the required face velocity from the Safetyflow fume cupboard specification (ILS Q0023979), or from AS/NZS 2243.8 where the specification does not give it - about 0.5 m/s average is typical and must be confirmed before use; (2) set the lowest single reading allowed; (3) confirm the monthly measurement frequency; (4) confirm the GM8902 anemometer and its calibration certificate are on site.
SWI-PRO-MET-206-004 | pH, DO and Conductivity Meter Calibration | Draft | Draft for review: Steps 3 and 4 follow the Martabe work instructions pH Meter TPS Cube Calibration (DOC-3-MET-PMC-WIN-00087-IE) and Operating pH Cube Meter (DOC-3-MET-PMC-WIN-00089-IE), and Steps 5 and 6 follow Operate DO Meter Portable (DOC-3-MET-PRS-WIN-00055-IE), v1.0, 25/12/2024. The TPS pH cube meter is adopted as at Martabe. pH calibration stays at three points (pH 4, 7 and 10) with the slope and offset recorded; the Martabe two-point calibration is not used. The Martabe DO method is applied to the Hanna HI9142 meter without the Oxyguard menu steps. JSEA-PRO-MET-206-004 must be updated for the added hazards before this instruction is approved.
SWI-PRO-MET-206-004 | pH, DO and Conductivity Meter Calibration | Open items | Open items: (1) confirm the TPS pH cube meter is on site, or procure it, and add it to the asset list; (2) take the DO calibration acceptance limit in air from the HI9142 manual - the Martabe value of 6 to 8 has no unit and is not used; (3) confirm from the TPS cube manual which control sets pH 4 - the source says "the same as pH 7"; (4) set the acceptable pH electrode slope range; (5) confirm the Mt. Morgan records for DO results and pH calibration - the Martabe Daily Task Checklist (DOC-FRM-00052-EN) and form DOC-4-MET-PMC-DFR-00142-EN are source names; (6) conductivity standards are not held - conductivity meters cannot be calibrated until they are.

## Keterangan
Register seri 206 Met Lab Operations and Safety (13 dokumen dalam format Heritage).
Data dibaca langsung dari Part 1 dan Part 2 tiap dokumen (judul, tanggal, owner, JSEA ID, area, frekuensi).
Klasifikasi dan padanan Martabe diambil dari docs/Matriks-Rujukan-Martabe.xlsx dan docs/Tabel-SWI-vs-Martabe.xlsx, ditimpa oleh keputusan review bila ada. Format Heritage tidak memuat kolom klasifikasi, jadi klasifikasi hanya ada di register.
Jumlah step dan halaman dihitung dari render dokumen (LibreOffice); halaman bisa sedikit berbeda di Microsoft Word.
Status isi: Final = tanpa tambahan Martabe terbuka; DRAFT FOR REVIEW = ada tambahan Martabe yang menunggu JSEA; Catatan = ada catatan yang perlu ditindaklanjuti.
Sheet "Draft dan open items" memuat paragraf draft, catatan, temuan safety-critical, dan open items di Description of work.