Safe Work Instruction

Cyanide Analyser Daily Operation and Check - 032-CA-001 and 051-CA-002

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots with sound tread - nitrile PVC boots at the sample point

- Chemical resistant suit, nitrile rubber gloves and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Harness and fall arrest where no fixed platform exists

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-207-007 - Cyanide Analyser Daily Operation and Check - 032-CA-001 and 051-CA-002 |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-207-007 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Operate the online cyanide analysers and carry out the daily check. Cyanide is dosed and recovered on what these analysers say. They also sit in a cyanide-bearing enclosure with their own sample pumps and leak detection. Replace the analyser filter sock (Step 3). Area: 207 Automatic Sampling and Online Analysers. Points and equipment covered: 032-CA-001 - CIL circuit online cyanide analyser (pH, free cyanide and WAD cyanide) 051-CA-002 - Cyanide adsorption (ReCYN) online cyanide analyser (pH, free cyanide and WAD cyanide) Frequency: To be set once the vendor manual is held. This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO211 - Automatic Sampling and Online Analyser Operation. Read that procedure before carrying out this task for the first time. Register line: SWI-PRO-MET-207-007 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-207-007. ISSUED BUT NOT FOR USE. The daily operation and check (Step 2) cannot be used until the named input arrives: the Molycop Cynoprobe v3 WAD vendor manual is not held and OS-040_B Cyanide Analyser is still open. The method, frequency and equipment are written when that input arrives. Draft for review: the filter sock replacement (Step 3) is taken from the Martabe work instruction Replace Filter Sock (DOC-3-MET-PMC-WIN-00091, undated), written for the Martabe analyser filter probe. JSEA-PRO-MET-207-007 must be updated for the added hazard before it is used. Open items: (1) confirm the equivalent filter probe and parts on 032-CA-001 and 051-CA-002, and whether these analysers have HCl dosing; (2) the source works from photographs that are not carried here - confirm the filtrate line arrangement for the flush; (3) the source gives no frequency for the filter sock change. This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Hydrogen cyanide gas released at the open sample point — Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Splash of cyanide-bearing slurry or solution to skin or eyes — Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. Contact with live electrical equipment at a sample pump, cutter or local control station — No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Inhalation of acid fume or reagent vapour at the bench — Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Non-representative sample leading to a wrong process decision — Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. [JSEA to confirm] HCl fume, slurry contact, stairs, slippery surfaces, dust and noise at the analyser filter probe — HCN alarm clear and the personal monitor on; HCl pump off with no fume before work. Latex gloves, dust mask, glasses, earplugs and safety boots. Eyewash proven flowing; wash contact with soap under running water. Chains and handrails in place on stairs with three points of contact. Keep the work area tidy and clear of slip hazards. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-207-007 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Working alone in a cyanide area - Gas monitor or fixed detection not healthy - Safety shower or eyewash not proven flowing - HCN alarm active at the analyser |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-207-007 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm a second person is in the area. Cyanide-area work is not done alone. - Take 5 and prepare 1 new filter sock, a 60 mm hose clamp, screwdriver, water hose, cutting knife and silicone. Check they are in good condition. - Notify the control room that the task is starting, and where you will be. - Confirm the personal gas monitor is on, in calibration date and reading clean. - Confirm the safety shower and eyewash nearest the work are within reach and proven flowing. - Confirm the area is fit - no restricted access, no gas alarm, no spill response in progress. Equipment Required : - To be confirmed once the Molycop Cynoprobe manual is held - New filter sock, 60 mm hose clamp, screwdriver, water hose, cutting knife, silicone - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Status - Not for Use - Task attempted without the vendor documentation |  | - This task cannot be completed until the Molycop Cynoprobe v3 WAD vendor manual is held. - The document is issued so the register line is not empty, and is marked NOT FOR USE. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Filter Sock Replacement - HCl fume at the filter probe - HCN in the cyanide-bearing enclosure - Slurry contact - Cuts from the cutting knife |  | - Switch the cyanide analyser off. Switch off the HCl pump and make sure no HCl fume is left. - Disconnect the pipe from the filter probe, then lift the bracket off the U bolt. - Wash the filter sock with water. - Remove the bolts holding the clamp and filter sock. Cut the old filter sock off with the cutting knife. - Flush the filtrate line: move the pump inlet line to the pump discharge line, and the pump discharge line into a bucket of water (arrangement to be confirmed on this analyser). - Fit the new filter sock. Join it to the filter cage, seal it with silicone and tighten it with the clamp. - Lift and dip the filter probe bracket back in place. - Switch the cyanide analyser on and check it works normally again. CAUTION: Work only with no HCN alarm active (alarm 5 ppm, high-high 10 ppm) and the personal HCN monitor on. The enclosure is cyanide-bearing - not done alone. Cut away from the body. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Spillage left on the ground - Equipment not returned - Control room not told the task is complete |  | - Leave the area clean. Recover spillage to the drain, not to the ground. - Return all equipment to its storage point. - Record what was done, what was found and anything not completed. - Notify the control room that the task is complete. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| HM-PRC-VXX-PRO211 | Automatic Sampling and Online Analyser Operation |
| JSEA-PRO-MET-207-007 | JSEA — Cyanide Analyser Daily Operation and Check - 032-CA-001 and 051-CA-002 |
| HM-MMM-WHS-REG-XXX-R00-SWI-SWMS | SWI / SWMS Register — tab SWI-MET |
| Note | No source held. Molycop Cynoprobe v3 WAD vendor manual not held; OS-040_B Cyanide Analyser still open. |
| DOC-3-MET-PMC-WIN-00091 | Martabe WI Replace Filter Sock (Indonesian, undated) - source of Step 3 |

In the event of an emergency

1. HCN Gas Alarm If the personal monitor or fixed detection alarms (5 ppm alarm, 10 ppm high-high), stop work, leave the area upwind and call CH19 “EMERGENCY”. Do not re-enter until the Shift Supervisor clears the area. 2. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 3. Skin or Eye Contact with Slurry Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 4. HCl Gas or Acid Contact Leave the area upwind. Flush skin or eyes with copious water at the safety shower or eyewash for a minimum of 15 minutes. If breathing is affected call CH19 “EMERGENCY” and call 000. 5. Contact with Live Electrical Equipment Do not touch the person. Switch off the supply at the nearest isolator if it is safe to do so, call CH19 “EMERGENCY” and call 000. 6. Slip, Fall or Lifting Injury Call CH19 “EMERGENCY”, call 000 if the injury is serious, and notify the Shift Supervisor.