Mt. Morgan SITE WORK INSTRUCTION

Resin Adsorption Capacity and Kinetics Test

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-205-012 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-205-012 |
| Area | 205 Met Lab - Chemical and Leach Testwork |
| Parent procedure | HM-PRC-VXX-PRO209 Metallurgical Laboratory Chemical and Leach Testwork |
| Classification | Operating |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

DRAFT FOR REVIEW - NOT FOR USE UNTIL APPROVED. The test method has been added from the Martabe work instructions 'Resin activity test' (03/03/2025) and 'Adsorption Test Fresh Resin' (09/01/2025). Acceptance targets for GG1200M and GG1200C still need vendor data - see the folder README. The source covers adsorption kinetics only; loading capacity is not covered by it.

# 1  TASK AND SCOPE

Measure the loading capacity and adsorption rate of the GG1200M and GG1200C resins.

Resin capacity is the design basis of the whole ReCYN circuit.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: To be set once the vendor data arrives.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO209 - Metallurgical Laboratory Chemical and Leach Testwork. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-205-012 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-205-012.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-205-012 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. |
| Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |
| Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. |
| [JSEA to confirm] Weighing and dissolving solid sodium cyanide and caustic; copper-cyanide solution at the bench | Weigh and dissolve inside the fume cupboard with airflow confirmed. Caustic added and pH verified above 10.5 BEFORE any cyanide is added. Full cyanide PPE per Section 4. Not done alone. Copper-cyanide waste to the cyanide-bearing stream only. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-205-012 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- Bottle roller

- Resin samples

- Screens 0.9 mm

- Solutions at known metal and cyanide concentration

- [M] Agitator with blade, baffle tank (1 L working volume)

- [M] Wet-sizing screens 1.18 mm and 1.4 mm [CONFIRM] against the 0.9 mm screen listed above

- [M] Balance, 250 mL sample bottles, filter paper, funnel, tea strainer, syringe, stopwatch, pH meter

- [M] NaCN, CuSO4.5H2O, NaOH (caustic), distilled water, fresh resin, loaded resin sample, Cu-eluted resin sample

- [M] KCN (KCN variant only)

- [M] Slurry from the CIL launder to detox, 10 g fresh resin per test, strainer to keep resin out of samples

# 6  STEPS

[M] = taken from the Martabe team work instruction "Resin activity test" (03/03/2025). [CONFIRM] = unspecified or inconsistent in the source; resolve before this instruction is approved for use.

Purpose [M]: compare the adsorption kinetics of each resin sample (loaded and Cu-eluted) against fresh resin.

Resin sizing [M]

1.  Pan the sample to separate carbon from resin.

2.  Stack the 1.18 mm and 1.4 mm screens.

3.  Wet-size the sample.

4.  Keep the resin that passes the 1.4 mm screen and is retained on the 1.18 mm screen. [CONFIRM] against the 0.9 mm screen in the equipment list.

Solution preparation [M]

5.  Check per cent solids. [CONFIRM] the sample this applies to is not stated in the source.

6.  Weigh 2 g NaCN.

7.  Measure 5 mL CuSO4.5H2O solution. [CONFIRM] solution strength is not stated.

8.  Put 5 L distilled water in the container and add the 5 mL copper sulphate.

9.  Hold the pH at 10 to 10.5 with caustic. [CONFIRM] this range reaches below 10.5; site rule in SWI-PRO-MET-205-007 is NEVER add cyanide below pH 10.5. Hold pH at 10.5 or above before the cyanide is added unless an approved deviation exists.

10.  Add the weighed NaCN and agitate until fully dissolved.

11.  Take a 20 mL head sample. Titrate for NaCN and submit to ITS for Cu assay.

Adsorption test [M]

12.  Put 1 L of the prepared solution in the baffle tank.

13.  Add 5 g of resin sample.

14.  At minute 5, 10, 15, 20 and 30 stop the agitator, take a solution sample of 30 mL, separate the resin from the solution and return the resin to the tank. [CONFIRM] the source text says 30 mL but the record table shows 20 mL per sample (1000 mL falling to 880 mL).

15.  Titrate each sample for cyanide.

16.  Send the remaining sample to ITS for Cu assay.

17.  Run fresh resin as the reference, then the loaded and Cu-eluted samples. Record on 'G:\Processing\5. Metallurgy\Metallurgy Lab\04 Metlab Data\Resin\5. Copper Loading Activity Test.xlsx' (columns: time, initial, sample and final volume, solution pH, NaCN mg/L).

KCN variant [M]: a later Martabe version, 'Resin activity test KCN' (dated '03/012/2025' - [CONFIRM] date), is identical except that it prepares 4 g KCN in place of 2 g NaCN. Its step text still says 'NaCN'. [CONFIRM] which reagent and mass is current: 4 g KCN is about 0.061 mol of cyanide against about 0.041 mol for 2 g NaCN (1.5 times), and 2.66 g KCN would be the molar equivalent. Do not run the KCN variant until this is settled.

Adsorption test of fresh resin on CIL slurry [M]

[M] = taken from the Martabe team work instruction "Adsorption Test Fresh Resin (file name 'Work_Instruction_Adsorption_Resin_-_2_gpl')" (09/01/2025). [CONFIRM] = unspecified or inconsistent in the source; resolve before this instruction is approved for use.

Purpose [M]: measure how well the resin adsorbs cyanide and copper from real slurry. The 10 g of resin in the source is 2 g/L, as the file name says.

Safety (added at review): the slurry is cyanide-bearing. Sampling and handling follow SWI-PRO-MET-201-004 and the HCN controls in Section 3, with the work in the fume cupboard. [CONFIRM] the source marks several lines as highlighted (the 150 mL sample volumes, the sample times, the filtering and the Cu and Fe assay) - confirm these are the approved values.

18.  Take 5 L of slurry (per resin type tested) from the launder of CIL tank 7 to detox. [CONFIRM] CIL 7 is a Martabe point; this plant's CIL train is 032-TK-002 to 032-TK-006. Name the equivalent sample point.

19.  Put the 5 L of slurry in the baffle tank. [CONFIRM] the source record table starts at 4000 mL initial volume, not 5 L.

20.  Take a 150 mL slurry sample as the 0 minute sample.

21.  Add 10 g of fresh resin to the baffle tank.

22.  Take 150 mL samples through a strainer so no resin is sampled, at minute 2, 5, 10, 20, 40 and 60. Record the pH at each time.

23.  Filter the solids from the solution.

24.  Titrate the filtered solution sample and send the remainder to ITS for NaCN titration and Cu and Fe assay.

25.  Enter the titration and pH results in the logsheet (columns: time, initial slurry volume, 150 mL sample, final volume, remaining solution, 106 mL solution sample, solution pH, NaCN mg/L).

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- KBK SOP-PROC-CNREC-01 and -12, received 11 Aug 2026 - operating loading and conditioning practice

- No Goldquip resin performance data held

- Martabe work instruction 'Resin activity test' (Indonesian, 03/03/2025) and its KCN version (dated 03/012/2025) - method source

- Martabe work instruction 'Adsorption Test Fresh Resin' (Indonesian, 09/01/2025; file name 'Work_Instruction_Adsorption_Resin_-_2_gpl') - source for the CIL slurry adsorption test