Safe Work Instruction

Slurry Viscosity and Rheology Test

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, nitrile rubber gloves and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-204-008 - Slurry Viscosity and Rheology Test |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-204-008 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Measure the viscosity and flow behaviour of a slurry at the operating density. Viscosity governs pumping duty, mixing and thickener behaviour, and rises sharply with fines. Daily viscosity is read on a Brookfield DV2TLV viscometer at 20 rpm, with a Marsh funnel quick check, following the Martabe method with the same instruments. Area: 204 Met Lab - Physical Testwork. Points and equipment covered: Met lab bench - physical testwork bench and instrument Frequency: On each new ore type, and whenever pumping or mixing behaviour changes. Draft for review: Steps 3 to 5 follow the Martabe work instructions Measure Viscosity Using Brookfield DV2TLV Viscometer (DOC-3-MET-PMC-WIN-00128-IE) and Measure Viscosity of Slurry Using Marsh Funnel Test (DOC-3-MET-PMC-WIN-00127-IE), v1.0, 25/12/2024. The Martabe method and instruments are adopted. JSEA-PRO-MET-204-008 must be updated for the added hazard before this instruction is approved. Open items: (1) the Marsh funnel source names the ASTM standard funnel but gives no fill volume, no sieving of coarse particles and no result in seconds per quart - set the fill volume and the reporting unit before use; (2) the Brookfield DV2T operating instruction and instrument manual are not received - confirm the spindle range for the plant slurries; (3) set the spindle and speed used for each slurry so all samples are read the same way. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Splash of cyanide-bearing slurry or solution to skin or eyes — Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. Contact with an agitator, sample cutter, pump or screen that starts without warning — Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. Manual handling injury from carrying sample buckets, containers and composite drums — Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. Non-representative sample leading to a wrong process decision — Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. [JSEA to confirm] Sample contact and splash to the eyes at the viscometer — Latex gloves, glasses and safety boots. Eyewash proven flowing; wash contact with soap under running water. Handle the sample carefully. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-204-008 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-204-008 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Take 5 and prepare the viscometer, spindles and measuring cup, and the Marsh funnel, bucket and stopwatch; check they work. Have clean water ready. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - Brookfield DV2TLV viscometer with spindles and its measuring cup - Marsh funnel, bucket and receiving container - Sample container and stand - Thermometer - Marcy scale - Stopwatch - Clean water for cleaning - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Sample Preparation - Sample not at the operating density - Settled slurry reading low - Splash of cyanide-bearing slurry |  | - Prepare the sample at the density the circuit actually runs at, not at a convenient density. - Record the density and temperature - both change the result substantially. - Homogenise the sample immediately before measuring; a settled slurry reads low. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Brookfield DV2TLV Set-up - Viscometer not level - Autozero with the spindle fitted - Spindle outside its range |  | - Make sure the viscometer is level. Switch it on with the power button. - Autozero by pressing next on the screen. The spindle must not be fitted during autozero. - Choose the spindle. Each spindle has a minimum and maximum viscosity it can measure; see the range in the manual. The spindle and speed chosen must be able to measure all the samples, and all samples must be measured with the same spindle and speed. - Fit the spindle. In the settings menu choose configure viscosity test. Set the speed between 1 and 200 rpm (20 rpm for daily work) and the measuring time for each sample. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Viscosity Measurement - Sample contact and splash to the eyes - Torque outside 10 to 90 per cent - Single point taken as a rheology curve |  | - Pour the sample into the cup. Bring the sample surface level with the mark on the spindle. Lower the spindle into the sample by adjusting the viscometer position. - Press run and record the viscosity reading. - If the torque is below 10 per cent or above 90 per cent, change the speed and the spindle. A slower speed and a smaller spindle can measure higher viscosity. - Where the test is used to size equipment, measure at the range of speeds the method calls for, record the full shear rate against shear stress series, and repeat at two other densities. CAUTION: Mineral slurries are not Newtonian. A single 20 rpm reading is for daily comparison only, not a rheology curve. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Marsh Funnel Check - Slurry spilled while pouring - Fill volume not the same each time |  | - Hold the funnel and close the bottom opening with a finger of the left hand. With the right hand pour the slurry from the bucket into the funnel. - Position the funnel over the receiving container. - Open the bottom opening and start the stopwatch at the same moment. Record the time for the sample to flow fully out of the funnel. - Stop the stopwatch when nothing more comes out of the funnel. CAUTION: Pour slurry slowly to avoid spills and splashes. A funnel time is a quick check, not a rheology curve. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Report and Instrument Clean-up - Viscosity reported without density and temperature - Spindle stored dirty |  | - Report the viscosity or the rheology curve, the density and the temperature together. A viscosity quoted without them means nothing. - Switch the viscometer off, take the spindle off, clean it with clean water and put it back in its box. - Clean the Marsh funnel and store it safely. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| JSEA-PRO-MET-204-008 | JSEA — Slurry Viscosity and Rheology Test |
| KBK-MIR-MP-PRO-OPE-SOP-0016 | Uji Viskositas |
| DOC-3-MET-PMC-WIN-00128-IE v1.0 | Martabe WI Measure Viscosity Using Brookfield DV2TLV Viscometer (25/12/2024) - source of Steps 3 and 4 |
| DOC-3-MET-PMC-WIN-00127-IE v1.0 | Martabe WI Measure Viscosity of Slurry Using Marsh Funnel Test (25/12/2024) - source of Step 5 |
| DOC-IV-MET-CHH-SOP-00039 | Martabe SOP referenced by the Martabe WI - not received |

In the event of an emergency

1. Contact with Moving Equipment If a person is caught by or struck by the viscometer or stirrer, hit the nearest emergency stop, call CH19 “EMERGENCY” and call 000. Do not approach the equipment until it is isolated and locked out. 2. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 3. Skin or Eye Contact with Slurry Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 4. Slip, Fall or Lifting Injury Call CH19 “EMERGENCY”, call 000 if the injury is serious, and notify the Shift Supervisor.