Safe Work Instruction

Fume Cupboard Operation, Scrubber Check and Cleaning

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots - nitrile PVC boots

- Chemical resistant suit, acid-resistant and nitrile rubber gloves, and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-206-002 - Fume Cupboard Operation, Scrubber Check and Cleaning |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-206-002 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Operate the fume cupboards and prove the scrubber is working before any fume-generating work. The fume cupboard and its scrubber are the only thing between the technician and the fume from acid digestion, cyanide work and perchloric acid. The fume cupboard face velocity is measured with the GM8902 anemometer, following the Martabe method, and accepted against the required face velocity. Area: 206 Met Lab - Facility, Safety and Quality. Points and equipment covered: Met lab bench - physical testwork bench and instrument Frequency: Every use; scrubber checked at the start of every shift; face velocity measured monthly and after any repair to the extraction fan, ductwork or scrubber. Draft for review: Steps 4 and 5 follow the Martabe work instruction Operating Anemometer GM8902 (DOC-3-MET-PMC-WIN-00130-IE, v1.0, 25/12/2024). The face velocity is measured at the marked working sash height, and the average and the lowest of 20 readings are recorded. JSEA-PRO-MET-206-002 must be updated for the added hazards before this instruction is approved. Open items: (1) take the required face velocity from the Safetyflow fume cupboard specification (ILS Q0023979), or from AS/NZS 2243.8 where the specification does not give it - about 0.5 m/s average is typical and must be confirmed before use; (2) set the lowest single reading allowed; (3) confirm the monthly measurement frequency; (4) confirm the GM8902 anemometer and its calibration certificate are on site. This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Inhalation of acid fume or reagent vapour at the bench — Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Contact with sulphuric or hydrochloric acid - burn to skin or eyes — Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. Hydrogen cyanide gas released at the open sample point — Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Cuts from broken glassware, and burns from hot glassware — Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. Cross-contamination from a container or cutter not cleaned between points — Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-206-002 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Fume cupboard or scrubber not running - Safety shower or eyewash not proven flowing - Instrument out of calibration |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-206-002 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm a second person is in the laboratory for any cyanide work. Cyanide work is not done alone. - Take 5 and prepare the anemometer, its data logger and a stopwatch where the face velocity is to be measured. Check the equipment is in good condition. - Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench. - Confirm the safety shower and eyewash are within reach and proven flowing. - Confirm the instrument or balance is within its calibration or verification interval. - Confirm the reagents to be used are in date and correctly labelled. - Inspect all glassware for chips and cracks and discard anything damaged. Equipment Required : - 2 x 2400 mm Safetyflow fume cupboards - Model A crossflow scrubber with pH dosing - Airflow indicator - pH meter for scrubber liquor - Washdown hose - Benetech GM8902 anemometer with its data logger, in calibration - Stopwatch - Log sheet - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Fume Cupboard and Scrubber Start-up - Extraction fan not running - Scrubber liquor out of its pH range - Fume released to the laboratory |  | - Confirm the extraction fan is running and the airflow indicator reads within its marked range before starting any work. - Confirm the scrubber is running and its pH dosing is operating, and record the scrubber liquor pH. - Where scrubber liquor pH is out of range, stop - a scrubber outside its pH range is not scrubbing. CAUTION: No acid or cyanide work at the bench until the airflow and the scrubber are proven. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Working in the Fume Cupboard - Sash raised above the working height - Rear baffle blocked - Perchlorate build-up in the ductwork |  | - Work with the sash at the marked working height. Raising the sash breaks the containment. - Keep apparatus at least 150 mm inside the sash line and never block the rear baffle. - For perchloric acid work, use only the cupboard rated for it, and carry out the washdown routine immediately after every use. - Wash down the working surface at the end of every session. CAUTION: Perchlorate build-up in ductwork is explosive. Washdown after every perchloric use. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Face Velocity Measurement - GM8902 Anemometer - Pinch point at the sash door - Breathing hazardous fume during the measurement - Anemometer out of calibration |  | - Make sure NOTHING is being done in the fume cupboard and the anemometer is calibrated. - Set the sash at the marked working height. Run the fume cupboard. - Plug the anemometer fan cable into the anemometer data logger. - Hold the anemometer vertically in the centre of the sash opening. Switch the anemometer on with the red button. - Record the reading every second, using the stopwatch, until you have 20 readings. CAUTION: Open and close the sash slowly and keep hands clear of the pinch point. Do not measure while any work is in the cupboard. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Acceptance and Records - Fume cupboard used below its required face velocity - Result not recorded |  | - Calculate the average of the 20 readings. Record the average and the lowest reading. - Compare the average with the required face velocity from the fume cupboard specification (see Open items). - Where the average is below the required face velocity, tag the fume cupboard out of service, do not use it, and report to the Met Lab supervisor. - Record the airflow reading, the scrubber pH, the face velocity and the washdown in the log. CAUTION: A fume cupboard below its required face velocity is not containing fume. Tag it out. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Cyanide-bearing residue to the general drain - Glassware left dirty - Result not traceable to the instrument |  | - Wash down the bench and the fume cupboard working surface. - Clean glassware and return it to the drying rack. - Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain. - Record the result, the method used and the instrument or balance identity. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| JSEA-PRO-MET-206-002 | JSEA — Fume Cupboard Operation, Scrubber Check and Cleaning |
| KBK-MIR-LAB-WET-SOP-054, -059 and ENV-SOP-061 | Fume cupboard and scrubber |
| ILS quotation Q0023979 | 2 x 2400 mm Safetyflow fume cupboard rated for perchloric and HF, Model A crossflow scrubber, exhaust fans |
| DOC-3-MET-PMC-WIN-00130-IE v1.0 | Martabe WI Operating Anemometer GM8902 (25/12/2024) - source of Step 4 |
| AS/NZS 2243.8 | Safety in laboratories - Fume cupboards (face velocity, see Open items) |

In the event of an emergency

1. HCN Gas Alarm If the personal monitor or fixed detection alarms (5 ppm alarm, 10 ppm high-high), stop work, leave the area upwind and call CH19 “EMERGENCY”. Do not re-enter until the Shift Supervisor clears the area. 2. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 3. Skin or Eye Contact with Reagent Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 4. HCl Gas or Acid Contact Leave the area upwind. Flush skin or eyes with copious water at the safety shower or eyewash for a minimum of 15 minutes. If breathing is affected call CH19 “EMERGENCY” and call 000. 5. Cuts from Broken Glassware Apply pressure to stop the bleeding, get a trained first aider, and report to the Shift Supervisor. Call CH19 “EMERGENCY” and 000 for a serious cut.