Mt. Morgan SITE WORK INSTRUCTION

Acid Spill Response and Neutralisation

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-206-006 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-206-006 |
| Area | 206 Met Lab - Facility, Safety and Quality |
| Parent procedure | HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality |
| Classification | Safety-critical |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  TASK AND SCOPE

Respond to an acid spill in the laboratory and neutralise it safely.

The laboratory holds concentrated sulphuric, hydrochloric and perchloric acid, and cyanide is in the same room.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Every acid handling task; drill at the set interval.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO210 - Metallurgical Laboratory Facility, Safety and Quality. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-206-006 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-206-006.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-206-006 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. |
| Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. |
| Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. |
| Splash of cyanide-bearing slurry or solution to skin or eyes | Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. |
| Cuts from broken glassware, and burns from hot glassware | Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-206-006 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- Hazchem spill kit

- Neutralising agent - soda ash or proprietary acid neutraliser

- Safety shower and eyewash

- Acid PPE

- Waste containers

- Fume cupboard

# 6  STEPS

1.  Contain the spill first, then neutralise. Never flush an acid spill to drain untreated.

2.  Where the spill is anywhere near cyanide-bearing material or waste, treat it as a potential HCN release - evacuate, raise the alarm and do not attempt to clean it up unprotected.

3.  Apply neutralising agent from the outside of the spill inwards, slowly. Adding neutraliser fast to a concentrated acid generates heat and spatter.

4.  Confirm neutralisation with pH paper before absorbing.

5.  Absorb, bag and label for chemical waste disposal.

6.  Ventilate the area and confirm the fume cupboard and general extraction are running.

7.  On skin or eye contact, flood at the safety shower or eyewash for at least 15 minutes and get medical attention. For hydrofluoric acid, apply calcium gluconate gel and treat as a medical emergency regardless of how the skin looks.

8.  Report and record every spill.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- KBK-MIR-LAB-WET-SOP-051 Acid spill handling

- ILS quotation Q0023979 - 120 L hazchem spill kit, safety shower and eye wash stations