Mt. Morgan SITE WORK INSTRUCTION

Ductless Fume Cabinet Filter Change and Breakthrough Check

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-206-013 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-206-013 |
| Area | 206 Met Lab - Facility, Safety and Quality |
| Parent procedure | HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality |
| Classification | Safety-critical |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  TASK AND SCOPE

Change the carbon filters on the ductless fume cabinet and check for breakthrough before it happens.

A ductless cabinet with a saturated filter discharges everything it has captured straight back into the room, and gives no indication that it is doing so.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Breakthrough check monthly; filter change on indication or at the manufacturer's interval.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO210 - Metallurgical Laboratory Facility, Safety and Quality. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-206-013 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-206-013.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-206-013 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. |
| Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |
| Manual handling injury from carrying sample buckets, containers and composite drums | Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-206-013 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- Purair 30-PP ductless cabinet

- Replacement carbon filters

- Airflow indicator

- Detector tubes or photoionisation detector for breakthrough

- Gloves and disposal bags

- Filter log

# 6  STEPS

1.  Confirm the cabinet airflow is within range before any use, and record it.

2.  Carry out the breakthrough check at the interval set - sample the discharge air with detector tubes or a suitable detector while a representative solvent or reagent is in use.

3.  Any detection in the discharge means the filter is spent. Take the cabinet out of service immediately and label it.

4.  Track filter life by usage, not by calendar alone. A cabinet used heavily saturates faster.

5.  To change: isolate the cabinet, wear gloves, remove the spent filter and bag it immediately as chemical waste.

6.  Fit the new filter, confirm the seal and re-check the airflow.

7.  Record the change date, the filter batch and the airflow reading in the filter log.

8.  Do not use the cabinet for any reagent the filter is not rated for. Carbon does not capture everything.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- ILS quotation Q0023979 - Purair 30-PP ductless cabinet with carbon filters

- No procedure was held for this cabinet before this revision