Safe Work Instruction

Cyanide Analyser Calibration and Standardisation

This document is the property of Heritage Minerals with all rights are reserved in respect of it. This document cannot be reproduced without the written consent of Heritage Minerals

# Purpose and Scope

This Safe Work Instruction is required so that the mechanical and operational integrity of the system is kept intact. It assists operators and mechanical technicians alike to identify possible problems before they happen.  It will help ensure that when personnel are performing this task it is performed in a safe and secure manner, to ensure the operational integrity of the machine and applies when this task is performed either by itself or in conjunction with other tasks in the area.

A JSA or Take 5 must be completed and discussed with the Supervisor if any variation to this SWI is necessary.

# RESPONSIBILITIES AND AUTHORITY

| What | Who | When |
| --- | --- | --- |
| Conform to procedure: Implement Changes: Change procedure: Review: Authorisation: | All personnel Anyone Training Facilitator Working party of content experts Manager | At all times As required As required On change & yearly On change |

# PPE REQUIREMENTS

- Safety helmet and safety glasses

- Chemical splash goggles and face shield

- High-visibility long-sleeved shirt and long trousers

- Safety boots with sound tread - nitrile PVC boots at the sample point

- Chemical resistant suit, nitrile rubber gloves and apron

- Personal HCN gas monitor, calibrated and in test date

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Harness and fall arrest where no fixed platform exists

In the event of an emergency

Call on CH19 “EMERGENCY, EMERGENCY, EMERGENCY” on the project to make everyone aware, call 000 for emergency services, police, fire or ambulance, you will need to stay calm and state when requested by a responder:

• Your Name. • Your Location. • Type of Emergency; and • Assistance Required.

| Part 1: Information |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Title: |  |  | SWI-PRO-MET-207-008 - Cyanide Analyser Calibration and Standardisation |  |  |  |  |  |  |  |  |  |  |  |
| Issued Date: |  |  | 11/08/2026 |  |  |  |  | Next Review Date: |  |  |  | 11/08/2027 |  |  |
| Document Owner: |  |  | Processing |  |  |  |  | JSEA ID: |  |  |  | JSEA-PRO-MET-207-008 |  |  |
| Part 2: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Description of work: |  |  |  |  |  | Calibrate and standardise the online cyanide analysers against known standards: the CIL circuit analyser 032-CA-001 (Steps 2 to 5) and the ReCYN tailing analyser 051-CA-002 (Step 6). Area: 207 Automatic Sampling and Online Analysers. Points and equipment covered: 032-CA-001 - CIL circuit online cyanide analyser (pH, free cyanide and WAD cyanide) 051-CA-002 - ReCYN tailing online cyanide analyser (pH, free cyanide and WAD cyanide) Frequency: To be set once the vendor calibration data is held. This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO211 - Automatic Sampling and Online Analyser Operation. Read that procedure before carrying out this task for the first time. Register line: SWI-PRO-MET-207-008 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-207-008. Draft for review: Steps 2 to 5 (032-CA-001) are taken from the Martabe work instruction Leach Analyzer Calibration (DOC-3-MET-PMC-WIN-00126-IE). Step 6 (051-CA-002) is taken from the Martabe Detox Analyser Calibration (DOC-3-MET-PMC-WIN-00117-IE): Mt. Morgan has no detox circuit - the ReCYN circuit follows CIL - so the Martabe detox analyser maps to the ReCYN tailing analyser 051-CA-002. Both v1.0, 25/12/2024. JSEA-PRO-MET-207-008 must be updated for the added hazards before this instruction is approved. Open items: (1) the source calibrates free NaCN only, with the WAD calibration box NOT ticked; the Mt. Morgan analysers read free and WAD cyanide, so the WAD calibration must be set from the Cynoprobe v3 WAD calibration data, which is not held; (2) standards for 032-CA-001 - the Indonesian source gives 250, 500 and 700 ppm NaCN, its English copy 30, 250 and 500 ppm; confirm; (3) standards for 051-CA-002 - the source materials list gives 30, 125 and 250 ppm, the steps give 30, 250 and 500 ppm; confirm against the ReCYN tailing range; (4) verification pH ranges in the source (9 to 11 for the leach analyser, 6 to 9 for the detox stream) reach below the 10.5 cyanide safety limit; confirm the range for each stream, in particular the ReCYN tailing; (5) the troubleshooting guide for critical errors is an attachment that was not received; (6) confirm the cyanide-bearing return point for leftover standard (the source says the leach tank) and the record folders. This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress. |  |  |  |  |  |  |  |  |
| Potential Hazards: |  |  |  |  |  | Hydrogen cyanide gas released at the open sample point — Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. Splash of cyanide-bearing slurry or solution to skin or eyes — Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. Contact with live electrical equipment at a sample pump, cutter or local control station — No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. Inhalation of acid fume or reagent vapour at the bench — Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. Cuts from broken glassware, and burns from hot glassware — Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. Non-representative sample leading to a wrong process decision — Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. [JSEA to confirm] Cyanide standards (2 L bottles) carried to the analyser hut; eye splash; stairs to the hut — Standards carried in lidded dark bottles inside a second plastic bag. Nitrile or latex gloves, glasses and the gas mask with A2B2E2K2P3 cartridge available. Safety data sheet read; eyewash proven flowing; spill kit at hand; pour slowly; clean spills at once. Poisoning or breathing difficulty: clinic or the emergency response team at once. Stairs with chains and handrails in place and three points of contact. [JSEA to confirm] Hand trapped at the ReCYN tailing analyser (051-CA-002) cover, and spilled standard to the skin or eyes — Latex gloves, glasses, gas mask and safety boots; keep hands clear of the cover pinch point and a safe distance. Spill kit at hand; clean spills at once with running water. HCN poisoning: clinic or the emergency response team at once. The full step-by-step hazard analysis, with the risk ranking, is in JSEA-PRO-MET-207-008 Section 4. |  |  |  |  |  |  |  |  |
| PERSONAL PROTECTIVE EQUIPMENT (PPE): |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Part 3: Safe Work Instruction Steps |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Job Number (Hazards/Cautions) |  | Application Method |  |  |  |  |  |  | Visual Reference |  |  |  |  |  |
| - Pre-start Check - Working alone in a cyanide area - Gas monitor or fixed detection not healthy - Safety shower or eyewash not proven flowing - Standards or materials out of date |  | - Confirm you are trained and signed off against this SWI, signed on to JSEA-PRO-MET-207-008 at a communication session, site inducted with current cyanide awareness and gas detection competency where the area requires it, and authorised by the Shift Supervisor. - Confirm a second person is in the area. Cyanide-area work is not done alone. - Take 5 and prepare the burette, flasks, pipette, dark bottles and bags; check they are in good condition. Confirm the materials are in date. If not, stop and report. - Prepare the potassium chloride electrolyte per KBK MET SOP-0053. - Notify the control room that the task is starting, and where you will be. - Confirm the personal gas monitor is on, in calibration date and reading clean. - Confirm the safety shower and eyewash nearest the work are within reach and proven flowing. - Confirm the area is fit - no restricted access, no gas alarm, no spill response in progress. Equipment Required : - Potassium chloride electrolyte per KBK MET SOP-0053 - NaCN standards at 3 levels, 2 L each, from the Mt. Morgan laboratory, for each analyser - AgNO3 0.01 M and rhodanine indicator - Burette, glass flasks, pipette, dark plastic bottles and plastic bags (second containment) - Fume-controlled position - Personal protective equipment as listed in PPE Requirements |  |  |  |  |  |  |  |  |  |  |  |  |
| - Standards and Titration - 032-CA-001 - Cyanide standard carried and spilled - Eye splash - HCN from the standards - Control room not aware of the calibration |  | - Take the NaCN standards at 3 levels, 2 L each, and put them in dark plastic bottles with good lids. Put each bottle in a plastic bag as second containment. - Confirm the nearest safety eyewash works. - Tell the control room the cyanide probe will be calibrated. The operator may take over the cyanide addition to the tank. - Titrate each standard by hand, 3 times per standard, and record the results. Use SWI-PRO-MET-205-001. - Calculate the average titrated concentration for each standard. CAUTION: Wear chemical-resistant gloves whenever you contact a cyanide standard. Carry the standards in lidded bottles inside a second bag. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Calibration Wizard - 032-CA-001 - Wrong calibration points entered - Odd reading kept in the calibration |  | - On the probe panel press setting. Select calibration with the down button and press enter. Press calibration wizard. - Use 3 calibration points and press continue. The source leaves the WAD calibration box NOT ticked - see open item (1) for the WAD calibration. - Set flushes between standards to 4, flushes between samples to 1 and sample times to 4. - Enter the average titrated concentration for each standard and press next. - Dip the auxiliary line (line 3) into standard 1 and press start, then press done. Repeat for the other two standards. Press next after all standards are measured. - Remove any reading that differs from the others, then press next. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Acceptance and Verification - 032-CA-001 - Calibration accepted with poor fit - pH reading below the cyanide safety limit |  | - Press accept if R squared is above 0.99. Record M and C before and after the calibration, and R squared after calibration. - Check the verification readings: automatic and manual NaCN within plus or minus 10 ppm; filtrate temperature 25 to 40 degrees C; analyser pH 9 to 11 (source range); probe pH immersed in the filtrate solution. If anything is abnormal, call the metallurgist. - Tell the control room the calibration is finished. Record the before and after readings in G:\Processing\5. Metallurgy\Metallurgy Lab\04 Metlab Data\Calibration Logsheet\Logsheet Cyanide Analyzer Calibration or on form DOC-4-MET-PMC-DFR-00136-EN (Form Cyanide Analyzer Calibration.xlsx). CAUTION: Cyanide solution below pH 10.5 releases HCN. Report any analyser pH reading below 10.5 to the metallurgist. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Standards Clean-up - 032-CA-001 - Leftover standard to the wrong drain |  | - Pour all the remaining standard solution into the cyanide-bearing return point (the source says the leach tank). Wash the bottles with water in the analyser hut and return them to storage. |  |  |  |  |  |  |  |  |  |  |  |  |
| - ReCYN Tailing Analyser Calibration - 051-CA-002 - Hand trapped at the analyser cover - Spilled standard to the skin or eyes - Calibration accepted with poor fit - pH reading below the cyanide safety limit |  | - Standards: 30, 250 and 500 ppm NaCN (see open item 3), 2 L each, from the laboratory, in dark lidded bottles inside a plastic bag. Confirm the eyewash at the ReCYN tailing analyser works and tell the control room the cyanide probe will be calibrated. Titrate each standard by hand 3 times (SWI-PRO-MET-205-001) and average the results. - Before calibrating, measure and record the automatic reading and the manual cyanide result. - Open the analyser cover. Calibrate as Steps 3 and 4: calibration wizard, 3 points, WAD box NOT ticked (see open item 1), flushes between standards 4, flushes between samples 1, sample times 4, enter the average titrated value for each standard, dip the auxiliary line in each standard, remove odd readings, accept if R squared is above 0.99, record M and C before and after and R squared after. - Open the cover of the analyser HMI screen. Check the automatic and manual NaCN agree within plus or minus 10 ppm; the filtrate temperature reads 25 to 40 degrees C; the analyser pH is in the range for the ReCYN tailing stream (source detox range 6 to 9 - see open item 4); the pH probe is immersed in the filtrate. If anything is abnormal, call the metallurgist. - Tell the control room the calibration is finished. Record the readings before and after in G:\Processing\5. Metallurgy\Metallurgy Lab\04 Metlab Data\Calibration Logsheet\Logsheet Detox Analyzer Calibration (source name) or on form DOC-4-MET-PMC-DFR-00138-EN (Form Detox Analyzer Calibration.xlsx). - Clean-up: pour all the remaining standard into the cyanide-bearing return point (the source says the leach tank). Wash the bottles with water and put them back in storage. CAUTION: Keep hands clear of the analyser cover pinch point. A ReCYN tailing stream below pH 10.5 can release HCN - keep the personal HCN monitor on. |  |  |  |  |  |  |  |  |  |  |  |  |
| - Completion and Clean Up - Spillage left on the ground - Equipment not returned - Control room not told the task is complete |  | - Leave the area clean. Recover spillage to the drain, not to the ground. - Return all equipment to its storage point. - Record what was done, what was found and anything not completed. - Notify the control room that the task is complete. - Record every value determined on the field or bench data sheet - date, time, shift and operator. Make the register or logbook entry and any handover signature. - Record any step not completed, with the reason. - Report to the Shift Supervisor, and record, any control in Part 2 found not in place. |  |  |  |  |  |  |  |  |  |  |  |  |

| Referenced Documents |  |
| --- | --- |
| Document ID | Document Title |
| HM-PRC-VXX-PRO211 | Automatic Sampling and Online Analyser Operation |
| JSEA-PRO-MET-207-008 | JSEA — Cyanide Analyser Calibration and Standardisation |
| HM-MMM-WHS-REG-XXX-R00-SWI-SWMS | SWI / SWMS Register — tab SWI-MET |
| KBK-MIR-MP-PRO-MET-SOP-0053 | Pembuatan Larutan Kalium Klorida untuk Cyanide Analizer, received 11 Aug 2026 - the electrolyte preparation only |
| Note | Molycop Cynoprobe v3 WAD calibration data not held |
| DOC-3-MET-PMC-WIN-00126-IE v1.0 | Martabe WI Leach Analyzer Calibration (25/12/2024) - source of Steps 2 to 5 (032-CA-001) |
| DOC-3-MET-PMC-WIN-00117-IE v1.0 | Martabe WI Detox Analyser Calibration (25/12/2024) - source of Step 6 (051-CA-002, ReCYN tailing) |
| DOC-2-MET-MEL-SOP-00087-IE and DOC-IV-MET-CHH-SOP-00039 | Martabe SOPs referenced by the Martabe WIs - not received |
| SWI-PRO-MET-205-001 | Free Cyanide Titration - Silver Nitrate Method |
| DOC-4-MET-PMC-DFR-00136-EN and 00138-EN | Form Cyanide Analyzer Calibration and Form Detox Analyzer Calibration |

In the event of an emergency

1. HCN Gas Alarm If the personal monitor or fixed detection alarms (5 ppm alarm, 10 ppm high-high), stop work, leave the area upwind and call CH19 “EMERGENCY”. Do not re-enter until the Shift Supervisor clears the area. 2. Suspected Cyanide Exposure If a person becomes dizzy, short of breath or collapses, call CH19 “EMERGENCY”, call 000, and bring the cyanide antidote kit and a trained first aider. Do not give mouth-to-mouth resuscitation. 3. Skin or Eye Contact with Slurry Flush with copious water at the safety shower or eyewash for a minimum of 15 minutes. Remove contaminated clothing while flushing. Report to the first aider. 4. Contact with Live Electrical Equipment Do not touch the person. Switch off the supply at the nearest isolator if it is safe to do so, call CH19 “EMERGENCY” and call 000. 5. Cuts from Broken Glassware Apply pressure to stop the bleeding, get a trained first aider, and report to the Shift Supervisor. Call CH19 “EMERGENCY” and 000 for a serious cut. 6. Slip, Fall or Lifting Injury Call CH19 “EMERGENCY”, call 000 if the injury is serious, and notify the Shift Supervisor.