Mt. Morgan SITE WORK INSTRUCTION

Bond Ball Mill Work Index Test

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-204-012 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-204-012 |
| Area | 204 Met Lab - Physical Testwork |
| Parent procedure | HM-PRC-VXX-PRO208 Metallurgical Laboratory Physical Testwork |
| Classification | Operating |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  TASK AND SCOPE

Determine the Bond ball mill work index of the ore.

The work index sets the milling energy the circuit needs. Getting it wrong sizes the mill wrong.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: On each new ore type or ore blend.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO208 - Metallurgical Laboratory Physical Testwork. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-204-012 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-204-012.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-204-012 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Contact with an agitator, sample cutter, pump or screen that starts without warning | Much of the plant is automated and equipment can start or stop without warning. Guards confirmed in place before approaching. Control room notified before the round starts and again on completion. Isolation and lock-out before any guard or cutter housing is opened. |
| Hearing damage in the mill, screen, compressor and blower areas | Hearing protection mandatory in signed areas. Double protection at the mill and compressor. Time at the point kept to what the task needs. |
| Inhalation of or contact with solid reagent - cyanide briquettes, lime or salt | Sampling of solid cyanide done only with the full cyanide PPE set including a gas mask with an A2B2E2K2P3 cartridge. Dust suppression and local extraction confirmed running. No dry sweeping. |
| Manual handling injury from carrying sample buckets, containers and composite drums | Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |
| Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-204-012 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- Bond ball mill and standard charge

- Sieve set

- Ro-Tap shaker

- Balance

- Riffle splitter

- Drying oven

# 6  STEPS

1.  Prepare the feed to the standard top size by stage crushing, not by prolonged grinding.

2.  Determine the feed size distribution and the F80.

3.  Charge the mill with the standard ball charge and the standard sample volume.

4.  Run the locked cycle test, screening at the closing screen and returning the oversize each cycle.

5.  Continue cycles until the circulating load stabilises at 250 per cent within the allowed tolerance.

6.  Determine the product size distribution and the P80 on the last cycles.

7.  Calculate the work index from the standard Bond equation.

8.  Report with the closing screen size, the F80 and the P80 - a work index without them is not usable.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- KBK-MIR-MP-PRO-OPE-SOP-0006 Uji Bond Work Index (BWI)