Mt. Morgan SITE WORK INSTRUCTION

Sieve Cleaning, Integrity Check and Aperture Verification

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-204-017 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-204-017 |
| Area | 204 Met Lab - Physical Testwork |
| Parent procedure | HM-PRC-VXX-PRO208 Metallurgical Laboratory Physical Testwork |
| Classification | Operating |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  TASK AND SCOPE

Keep the sieve set clean, intact and verified, so a sizing result means what it says.

A blinded or torn sieve produces a wrong size distribution that looks perfectly normal on the sheet.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: After every use for cleaning; quarterly for integrity and aperture verification.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO208 - Metallurgical Laboratory Physical Testwork. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-204-017 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-204-017.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-204-017 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Inhalation of or contact with solid reagent - cyanide briquettes, lime or salt | Sampling of solid cyanide done only with the full cyanide PPE set including a gas mask with an A2B2E2K2P3 cartridge. Dust suppression and local extraction confirmed running. No dry sweeping. |
| Hearing damage in the mill, screen, compressor and blower areas | Hearing protection mandatory in signed areas. Double protection at the mill and compressor. Time at the point kept to what the task needs. |
| Cuts from broken glassware, and burns from hot glassware | Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |
| Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-204-017 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- Ultrasonic sieve cleaner FSS Sieve 200 DH

- Soft brushes

- Compressed air at low pressure

- Magnifier or microscope

- Sieve register

- Certified reference sieves or a check sample

# 6  STEPS

1.  Clean each sieve immediately after use - brush from the underside, never force material through from above.

2.  Use the ultrasonic cleaner for blinded fine sieves. Do not use a wire brush on any sieve.

3.  Inspect the mesh under magnification for tears, stretched apertures, and separation from the frame.

4.  Withdraw and mark any damaged sieve immediately. A damaged sieve left in the rack will be used.

5.  Verify apertures quarterly against certified reference sieves or by running a check sample of known distribution.

6.  Record every clean, inspection and verification in the sieve register against the sieve identity.

7.  Replace sieves on the interval the register sets, or on failure, whichever comes first.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- ILS quotation Q0023979 - ultrasonic sieve cleaner FSS Sieve 200 DH

- No cleaning or aperture verification procedure was held before this revision