Mt. Morgan SITE WORK INSTRUCTION

Free Cyanide Titration - Silver Nitrate Method

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-205-001 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-205-001 |
| Area | 205 Met Lab - Chemical and Leach Testwork |
| Parent procedure | HM-PRC-VXX-PRO209 Metallurgical Laboratory Chemical and Leach Testwork |
| Classification | Safety-critical |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  TASK AND SCOPE

Determine free cyanide by titration against standardised silver nitrate.

Free cyanide is the single most important solution number on this plant - it drives dosing, recovery and the environmental obligation.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Every shift on every cyanide-bearing stream on the schedule.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO209 - Metallurgical Laboratory Chemical and Leach Testwork. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-205-001 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-205-001.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-205-001 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. |
| Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. |
| Cuts from broken glassware, and burns from hot glassware | Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. |
| Splash of cyanide-bearing slurry or solution to skin or eyes | Sample valve opened slowly with the container positioned first and the body out of the spray path. Face shield, chemical splash goggles, chemical-resistant gloves and apron. Never open a sample valve whose discharge cannot be seen. Safety shower and eyewash proven flowing before the round starts. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |
| Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-205-001 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- Burette 50 mL and stand

- Pipette 10 or 25 mL

- Conical flasks 250 mL

- Standardised AgNO3 0.01 M

- Rhodanine indicator

- Wash bottle and distilled water

- Fume cupboard

# 6  STEPS

1.  Work in the fume cupboard with the airflow confirmed. Cyanide solutions release HCN, faster if the pH has dropped.

2.  Filter the sample if it is not already clear, and titrate as soon after filtration as possible to limit volatilisation loss.

3.  Rinse the burette with distilled water then condition it three times with standardised AgNO3.

4.  Fill the burette and expel air from the tip.

5.  Pipette the aliquot into a conical flask.

6.  Add 3 to 5 drops of rhodanine indicator - the solution turns pale yellow.

7.  Add AgNO3 dropwise with continuous swirling. As the endpoint nears a salmon pink precipitate forms and redissolves slowly; slow the addition there.

8.  Continue dropwise until a permanent salmon pink holds for 30 seconds without swirling.

9.  Record the burette reading and calculate free cyanide.

10.  Run a duplicate on every tenth sample and a standard once per shift.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- KBK-MIR-MP-PRO-OPE-SOP-0010 Titrasi Sianida

- KBK-MIR-MP-PRO-MET-SOP-0037 Titrasi Sianida dengan Perak Nitrat and SOP-0043 AgNO3 0.01 M standard, both received 11 Aug 2026