Mt. Morgan SITE WORK INSTRUCTION

Automatic Titrator Operation and Method Set-Up

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-205-018 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-205-018 |
| Area | 205 Met Lab - Chemical and Leach Testwork |
| Parent procedure | HM-PRC-VXX-PRO209 Metallurgical Laboratory Chemical and Leach Testwork |
| Classification | Operating |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  TASK AND SCOPE

Set up and run the Hanna HI902C automatic titrator with the methods the plant needs.

The titrator replaces manual titration for routine work. A wrongly configured method produces confident wrong answers all shift.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: At set-up, and whenever a method is added or changed.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO209 - Metallurgical Laboratory Chemical and Leach Testwork. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-205-018 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-205-018.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-205-018 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. |
| Contact with sodium hydroxide solution - burn to skin or eyes | Full chemical PPE including face shield and chemical-resistant gloves. Sample drawn slowly from the designated valve. Safety shower and eyewash within reach and proven flowing. |
| Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. |
| Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. |
| Cuts from broken glassware, and burns from hot glassware | Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. |
| Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |
| Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-205-018 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- Hanna HI902C pH/ORP/ISE titrator

- Electrodes and ISE probes

- Standard solutions

- Titrant reservoirs

- Beakers and stirrers

- Fume cupboard

# 6  STEPS

1.  Confirm the correct electrode or ISE is fitted for the determination and that it is conditioned.

2.  Load the titrant and prime the burette, expelling all air.

3.  Enter the method parameters from the manual method being ported - aliquot, titrant concentration, endpoint criterion, stirring and dosing rate.

4.  Run the method against a known standard first. Do not run samples until the standard reads correctly.

5.  Compare the titrator result against a manual titration on the same sample, in duplicate.

6.  Where the two disagree beyond tolerance, correct the method - do not adjust the result.

7.  Save the method under a controlled name and version and record it in the method register.

8.  Re-verify against a manual titration weekly and after any electrode change.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- ILS quotation Q0023979 - Hanna HI902C automatic titrator with ISE

- KBK-MIR-MP-PRO-MET-SOP-0025, -0037, -0039, -0040, -0041, -0042 and -0043 - the seven titration methods to port, all received 11 Aug 2026