Mt. Morgan SITE WORK INSTRUCTION

Automatic Titrator Electrode Care, Standardisation and Verification

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-205-019 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-205-019 |
| Area | 205 Met Lab - Chemical and Leach Testwork |
| Parent procedure | HM-PRC-VXX-PRO209 Metallurgical Laboratory Chemical and Leach Testwork |
| Classification | Operating |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  TASK AND SCOPE

Keep the titrator electrodes in condition and prove they are reading correctly.

Electrodes drift and age. An unverified electrode is the most common cause of a titrator giving wrong results that nobody questions.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Daily standardisation; weekly verification; replacement on failure.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO209 - Metallurgical Laboratory Chemical and Leach Testwork. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-205-019 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-205-019.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-205-019 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. |
| Cuts from broken glassware, and burns from hot glassware | Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. |
| Contact with live electrical equipment at a sample pump, cutter or local control station | No opening of electrical enclosures. Faults reported, not investigated. Isolation, lock-out, tag-out and test-for-dead by an authorised person before any electrical work. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |
| Non-representative sample leading to a wrong process decision | Line flushed for the stated time before collecting. Sample taken at the same point, in the same way and at the same frequency every time. Container labelled BEFORE the sample is taken, with point, date, time and shift. Steady-state confirmed for 15 to 20 minutes before sampling. A sample taken off-condition is marked as such or discarded. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-205-019 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- pH buffers 4, 7 and 10

- ISE standards

- Electrode storage solution

- Distilled water

- Soft tissue

- Electrode log

# 6  STEPS

1.  Standardise the pH electrode daily against buffers 4, 7 and 10, and record the slope and offset.

2.  Reject the electrode where the slope falls outside the acceptable range - do not keep using it because it still gives a number.

3.  Standardise the ISE against its standards at the interval the method sets.

4.  Rinse electrodes with distilled water between samples and blot dry - never wipe, which strips the junction.

5.  Store electrodes in the correct storage solution. Never store in distilled water; it leaches the reference.

6.  Check and top up the reference electrolyte at the interval the manual sets.

7.  Verify against an independent standard weekly and record it in the electrode log.

8.  Record every standardisation, verification and replacement against the electrode serial number.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- KBK-MIR-MP-PRO-MET-SOP-0030 Kalibrasi DO-meter and SOP-0031 Kalibrasi pH-meter, both received 11 Aug 2026

- KBK-MIR-MP-PRO-MET-SOP-0043 AgNO3 standard, received 11 Aug 2026

- ILS quotation Q0023979