Mt. Morgan SITE WORK INSTRUCTION

Met Lab Chemical Receipt, Storage and Decanting

Standard Work Instruction  ·  Heritage Minerals  ·  Mt. Morgan Processing Plant

| Document number | SWI-PRO-MET-206-001 |
| --- | --- |
| JSEA reference | JSEA-PRO-MET-206-001 |
| Area | 206 Met Lab - Facility, Safety and Quality |
| Parent procedure | HM-PRC-VXX-PRO210 Metallurgical Laboratory Facility, Safety and Quality |
| Classification | Safety-critical |
| Revision | Rev A |
| Date | 11 August 2026 |
| Prepared by | GGT (cost code 4034-033) |
| Reviewed by | ________________________ |
| Approved by | ________________________ |

# 1  TASK AND SCOPE

Receive, store and decant laboratory chemicals safely and in a way that keeps them fit for use.

The laboratory holds concentrated acids, cyanide and oxidisers in one room. Segregation and labelling are the controls.

Points and equipment covered:

| Tag / point | Description |
| --- | --- |
| Met lab bench | Physical testwork bench and instrument |

Frequency: Every delivery, and every decant.

This instruction sits under Standard Work Procedure HM-PRC-VXX-PRO210 - Metallurgical Laboratory Facility, Safety and Quality. Read that procedure before carrying out this task for the first time.

Register line: SWI-PRO-MET-206-001 in HM-MMM-WHS-REG-XXX-R00-SWI-SWMS, tab SWI-MET. The matching hazard analysis is JSEA-PRO-MET-206-001.

# 2  WHO MAY DO THIS

- Trained and signed off against this instruction.

- Signed on to JSEA JSEA-PRO-MET-206-001 at a communication session.

- Site inducted, with current cyanide awareness and gas detection competency where the area requires it.

- Authorised by the Shift Supervisor before the task starts.

This is a safety-critical task. It is not carried out alone in a cyanide area, and not without the Shift Supervisor knowing it is in progress.

# 3  HAZARDS SPECIFIC TO THIS TASK

| Hazard | Control |
| --- | --- |
| Contact with sulphuric or hydrochloric acid - burn to skin or eyes | Full acid PPE including face shield, acid-resistant gloves, apron and boots. Sample drawn only from the designated sample valve. Add acid to water, never the reverse. Safety shower and eyewash within reach and proven flowing. Spill kit at the point. |
| Contact with sodium hydroxide solution - burn to skin or eyes | Full chemical PPE including face shield and chemical-resistant gloves. Sample drawn slowly from the designated valve. Safety shower and eyewash within reach and proven flowing. |
| Hydrogen cyanide gas released at the open sample point | Personal HCN monitor worn and switched on, in calibration date. Fixed detection confirmed healthy - alarm 5 ppm, high-high 10 ppm. Sample point approached from upwind. pH of the stream confirmed above 10.5 before opening. Second person in sight and clear. Cyanide antidote kit on site and a trained first aider on shift. Withdraw upwind and report on any reading above the site action level. |
| Inhalation of acid fume or reagent vapour at the bench | Work carried out inside the fume cupboard with the sash at the marked working height and the airflow confirmed before starting. Scrubber running and its dosing confirmed. Perchloric work only in the cupboard rated for it, with the washdown routine after every use. No work at an open bench. |
| Cuts from broken glassware, and burns from hot glassware | Glassware inspected for chips and cracks before use and discarded if damaged. Hot glassware handled with tongs or heat-resistant gloves and set down on a heat mat. Broken glass to the sharps bin, never to the general waste or by hand. |
| Manual handling injury from carrying sample buckets, containers and composite drums | Sample volume kept to the schedule - no over-filling. Two-person lift or a trolley for composite containers above the site manual handling limit. Route cleared before lifting. |
| Cross-contamination from a container or cutter not cleaned between points | Bucket, container, cutter and sieve rinsed with process water and then with the stream being sampled, and the rinse discarded, before every sample. Dedicated containers for high-grade streams. Labels checked against the schedule before handover. |

The full step-by-step hazard analysis, with the risk ranking, is in JSEA JSEA-PRO-MET-206-001 Section 4.

# 4  PPE

- Hard hat

- Safety glasses

- High-visibility long-sleeved shirt and long trousers

- Safety boots

- Gloves appropriate to the task

- Personal HCN gas monitor, calibrated and in test date

- Chemical splash goggles and face shield

- Nitrile rubber gloves

- Nitrile PVC boots

- Chemical resistant suit

- Full or half gas mask with A2B2E2K2P3 cartridge, carried and available for immediate use

- Acid-resistant gloves, apron and face shield

- Chemical splash goggles

# 5  BEFORE YOU START

Every item below must be true. If one is not, stop and report it.

- Confirm the fume cupboard airflow, the scrubber and its dosing are running before any acid or cyanide work at the bench.

- Confirm the safety shower and eyewash are within reach and proven flowing.

- Confirm the instrument or balance is within its calibration or verification interval.

- Confirm the reagents to be used are in date and correctly labelled.

- Inspect all glassware for chips and cracks and discard anything damaged.

Equipment required:

- 160 L flammable, 160 L corrosive and 60 L toxic cabinets

- Fume cupboard

- Decanting equipment

- Spill kit

- Chemical register

- Labels

- Safety data sheets

# 6  STEPS

1.  Check the delivery against the order and the safety data sheet before accepting it.

2.  Reject anything with a damaged container, an illegible label or no safety data sheet.

3.  Log the chemical into the register with quantity, batch, date received and expiry.

4.  Store to the correct cabinet. Acids away from cyanide, oxidisers away from flammables, and nothing on the floor.

5.  NEVER store an acid where a leak could reach cyanide. Acid on cyanide releases hydrogen cyanide gas.

6.  Decant only in the fume cupboard, with the airflow confirmed and a spill kit at hand.

7.  Label every decanted container immediately with contents, concentration, date and preparer.

8.  Return the parent container to its cabinet before leaving the bench.

# 7  ON COMPLETION

- Wash down the bench and the fume cupboard working surface.

- Clean glassware and return it to the drying rack.

- Dispose of residues to the correct waste stream - cyanide-bearing effluent does NOT go to the general drain.

- Record the result, the method used and the instrument or balance identity.

# 8  RECORDS

- Field or bench data sheet - date, time, shift, operator, and every value determined.

- Register or logbook entry and any handover signature.

- Any step not completed, recorded with the reason.

- Any control in Section 3 found not in place, reported to the Shift Supervisor and recorded.

# 9  REFERENCE

- KBK-MIR-LAB-WET-SOP-025 chemical issue and WET-SOP-047 acid transfer

- ILS quotation Q0023979 - 160 L flammable, 160 L corrosive and 60 L toxic/poisons cabinets